## Info
Material | Density (g/cc) | Relative Electron Densiy | Effective Atomic Number
Water | 1 | 1 | 7.417
Air | 0.0013 | 0.0012 | 8
Acrylic | 1.18 | 1.147 | 6.467
Bone20 | 1.14 | 1.084 | 9.089
Bone50 | 1.4 | 1.312 | 11.455
Delrin | 1.42 | 1.363 | 6.95
LDPE | 0.92 | 0.945 | 5.444
PMP | 0.83 | 0.853 | 5.444
Polystyrene | 1.03 | 0.998 | 5.695
Teflon | 2.16 | 1.868 | 8.433
Lung | 0.19 | 0.297 | 6.642
On the following worksheets are mass attenuation coefficients for sensitometry materials used in Catphan® phantoms. Data is provided for selected energies from 20 keV to 20 MeV. Contributions from different interactions are given as well as totals both with and without coherent scattering effects. The values were obtained from the NIST XCOM database using our best knowledge of material compositions. The data is subject to change pending new information.
2013-03-27

## Air
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.08649 | 0.1558 | 0.5077 | 0 | 0 | 0.75 | 0.6635
0.021 | 0.08041 | 0.157 | 0.4341 | 0 | 0 | 0.6715 | 0.5911
0.022 | 0.07493 | 0.158 | 0.3738 | 0 | 0 | 0.6067 | 0.5318
0.023 | 0.06997 | 0.1589 | 0.324 | 0 | 0 | 0.5529 | 0.4829
0.024 | 0.06547 | 0.1597 | 0.2825 | 0 | 0 | 0.5076 | 0.4422
0.025 | 0.06138 | 0.1604 | 0.2476 | 0 | 0 | 0.4694 | 0.408
0.026 | 0.05766 | 0.161 | 0.2181 | 0 | 0 | 0.4368 | 0.3791
0.027 | 0.05425 | 0.1615 | 0.1931 | 0 | 0 | 0.4088 | 0.3546
0.028 | 0.05113 | 0.1619 | 0.1716 | 0 | 0 | 0.3847 | 0.3336
0.029 | 0.04826 | 0.1623 | 0.1532 | 0 | 0 | 0.3638 | 0.3155
0.03 | 0.04563 | 0.1626 | 0.1372 | 0 | 0 | 0.3455 | 0.2999
0.031 | 0.0432 | 0.1629 | 0.1234 | 0 | 0 | 0.3295 | 0.2863
0.032 | 0.04096 | 0.1631 | 0.1113 | 0 | 0 | 0.3153 | 0.2744
0.033 | 0.03889 | 0.1632 | 0.1007 | 0 | 0 | 0.3028 | 0.2639
0.034 | 0.03696 | 0.1633 | 0.09138 | 0 | 0 | 0.2917 | 0.2547
0.035 | 0.03518 | 0.1634 | 0.08316 | 0 | 0 | 0.2817 | 0.2466
0.036 | 0.03352 | 0.1634 | 0.07587 | 0 | 0 | 0.2728 | 0.2393
0.037 | 0.03197 | 0.1634 | 0.06939 | 0 | 0 | 0.2648 | 0.2328
0.038 | 0.03053 | 0.1634 | 0.06361 | 0 | 0 | 0.2575 | 0.227
0.039 | 0.02918 | 0.1633 | 0.05845 | 0 | 0 | 0.251 | 0.2218
0.04 | 0.02792 | 0.1633 | 0.05381 | 0 | 0 | 0.245 | 0.2171
0.041 | 0.02674 | 0.1632 | 0.04965 | 0 | 0 | 0.2395 | 0.2128
0.042 | 0.02564 | 0.163 | 0.04589 | 0 | 0 | 0.2346 | 0.2089
0.043 | 0.0246 | 0.1629 | 0.04249 | 0 | 0 | 0.23 | 0.2054
0.044 | 0.02362 | 0.1627 | 0.03942 | 0 | 0 | 0.2258 | 0.2021
0.045 | 0.0227 | 0.1625 | 0.03663 | 0 | 0 | 0.2219 | 0.1992
0.046 | 0.02183 | 0.1623 | 0.03409 | 0 | 0 | 0.2183 | 0.1964
0.047 | 0.02101 | 0.1621 | 0.03177 | 0 | 0 | 0.2149 | 0.1939
0.048 | 0.02024 | 0.1619 | 0.02965 | 0 | 0 | 0.2118 | 0.1916
0.049 | 0.0195 | 0.1617 | 0.02772 | 0 | 0 | 0.2089 | 0.1894
0.05 | 0.01881 | 0.1614 | 0.02594 | 0 | 0 | 0.2062 | 0.1874
0.051 | 0.01815 | 0.1612 | 0.02431 | 0 | 0 | 0.2037 | 0.1855
0.052 | 0.01753 | 0.1609 | 0.02281 | 0 | 0 | 0.2013 | 0.1838
0.053 | 0.01694 | 0.1607 | 0.02143 | 0 | 0 | 0.1991 | 0.1821
0.054 | 0.01638 | 0.1604 | 0.02016 | 0 | 0 | 0.1969 | 0.1806
0.055 | 0.01584 | 0.1601 | 0.01898 | 0 | 0 | 0.195 | 0.1791
0.056 | 0.01533 | 0.1599 | 0.01789 | 0 | 0 | 0.1931 | 0.1777
0.057 | 0.01484 | 0.1596 | 0.01688 | 0 | 0 | 0.1913 | 0.1765
0.058 | 0.01438 | 0.1593 | 0.01595 | 0 | 0 | 0.1896 | 0.1752
0.059 | 0.01394 | 0.159 | 0.01508 | 0 | 0 | 0.188 | 0.1741
0.06 | 0.01352 | 0.1587 | 0.01427 | 0 | 0 | 0.1865 | 0.173
0.061 | 0.01311 | 0.1584 | 0.01351 | 0 | 0 | 0.185 | 0.1719
0.062 | 0.01273 | 0.1581 | 0.01281 | 0 | 0 | 0.1836 | 0.1709
0.063 | 0.01236 | 0.1578 | 0.01216 | 0 | 0 | 0.1823 | 0.1699
0.064 | 0.01201 | 0.1575 | 0.01154 | 0 | 0 | 0.181 | 0.169
0.065 | 0.01167 | 0.1572 | 0.01097 | 0 | 0 | 0.1798 | 0.1681
0.066 | 0.01134 | 0.1568 | 0.01043 | 0 | 0 | 0.1786 | 0.1673
0.067 | 0.01103 | 0.1565 | 0.009932 | 0 | 0 | 0.1775 | 0.1665
0.068 | 0.01073 | 0.1562 | 0.00946 | 0 | 0 | 0.1764 | 0.1657
0.069 | 0.01045 | 0.1559 | 0.009018 | 0 | 0 | 0.1754 | 0.1649
0.07 | 0.01017 | 0.1556 | 0.008601 | 0 | 0 | 0.1744 | 0.1642
0.071 | 0.009905 | 0.1553 | 0.00821 | 0 | 0 | 0.1734 | 0.1635
0.072 | 0.00965 | 0.1549 | 0.007842 | 0 | 0 | 0.1724 | 0.1628
0.073 | 0.009405 | 0.1546 | 0.007494 | 0 | 0 | 0.1715 | 0.1621
0.074 | 0.009169 | 0.1543 | 0.007167 | 0 | 0 | 0.1706 | 0.1615
0.075 | 0.008942 | 0.154 | 0.006858 | 0 | 0 | 0.1698 | 0.1608
0.076 | 0.008722 | 0.1537 | 0.006566 | 0 | 0 | 0.1689 | 0.1602
0.077 | 0.008511 | 0.1533 | 0.00629 | 0 | 0 | 0.1681 | 0.1596
0.078 | 0.008307 | 0.153 | 0.006029 | 0 | 0 | 0.1674 | 0.159
0.079 | 0.008111 | 0.1527 | 0.005782 | 0 | 0 | 0.1666 | 0.1585
0.08 | 0.007921 | 0.1524 | 0.005548 | 0 | 0 | 0.1658 | 0.1579
0.081 | 0.007738 | 0.1521 | 0.005326 | 0 | 0 | 0.1651 | 0.1574
0.082 | 0.007561 | 0.1517 | 0.005116 | 0 | 0 | 0.1644 | 0.1568
0.083 | 0.007389 | 0.1514 | 0.004916 | 0 | 0 | 0.1637 | 0.1563
0.084 | 0.007224 | 0.1511 | 0.004727 | 0 | 0 | 0.163 | 0.1558
0.085 | 0.007064 | 0.1508 | 0.004547 | 0 | 0 | 0.1624 | 0.1553
0.086 | 0.006909 | 0.1504 | 0.004375 | 0 | 0 | 0.1617 | 0.1548
0.087 | 0.006759 | 0.1501 | 0.004213 | 0 | 0 | 0.1611 | 0.1543
0.088 | 0.006614 | 0.1498 | 0.004057 | 0 | 0 | 0.1605 | 0.1539
0.089 | 0.006474 | 0.1495 | 0.00391 | 0 | 0 | 0.1599 | 0.1534
0.09 | 0.006338 | 0.1492 | 0.003769 | 0 | 0 | 0.1593 | 0.1529
0.091 | 0.006206 | 0.1489 | 0.003635 | 0 | 0 | 0.1587 | 0.1525
0.092 | 0.006078 | 0.1485 | 0.003507 | 0 | 0 | 0.1581 | 0.1521
0.093 | 0.005954 | 0.1482 | 0.003384 | 0 | 0 | 0.1576 | 0.1516
0.094 | 0.005834 | 0.1479 | 0.003268 | 0 | 0 | 0.157 | 0.1512
0.095 | 0.005717 | 0.1476 | 0.003156 | 0 | 0 | 0.1565 | 0.1508
0.096 | 0.005604 | 0.1473 | 0.00305 | 0 | 0 | 0.156 | 0.1503
0.097 | 0.005494 | 0.147 | 0.002948 | 0 | 0 | 0.1554 | 0.1499
0.098 | 0.005387 | 0.1467 | 0.00285 | 0 | 0 | 0.1549 | 0.1495
0.099 | 0.005284 | 0.1464 | 0.002757 | 0 | 0 | 0.1544 | 0.1491
0.1 | 0.005183 | 0.1461 | 0.002667 | 0 | 0 | 0.1539 | 0.1487
0.101 | 0.005085 | 0.1458 | 0.002582 | 0 | 0 | 0.1534 | 0.1483
0.102 | 0.00499 | 0.1455 | 0.0025 | 0 | 0 | 0.153 | 0.148
0.103 | 0.004897 | 0.1452 | 0.002421 | 0 | 0 | 0.1525 | 0.1476
0.104 | 0.004807 | 0.1449 | 0.002346 | 0 | 0 | 0.152 | 0.1472
0.105 | 0.00472 | 0.1446 | 0.002273 | 0 | 0 | 0.1516 | 0.1468
0.106 | 0.004635 | 0.1443 | 0.002204 | 0 | 0 | 0.1511 | 0.1465
0.107 | 0.004552 | 0.144 | 0.002137 | 0 | 0 | 0.1507 | 0.1461
0.108 | 0.004471 | 0.1437 | 0.002073 | 0 | 0 | 0.1502 | 0.1457
0.109 | 0.004392 | 0.1434 | 0.002011 | 0 | 0 | 0.1498 | 0.1454
0.11 | 0.004316 | 0.1431 | 0.001952 | 0 | 0 | 0.1494 | 0.145
0.111 | 0.004241 | 0.1428 | 0.001895 | 0 | 0 | 0.1489 | 0.1447
0.112 | 0.004168 | 0.1425 | 0.00184 | 0 | 0 | 0.1485 | 0.1444
0.113 | 0.004098 | 0.1422 | 0.001787 | 0 | 0 | 0.1481 | 0.144
0.114 | 0.004028 | 0.1419 | 0.001737 | 0 | 0 | 0.1477 | 0.1437
0.115 | 0.003961 | 0.1417 | 0.001688 | 0 | 0 | 0.1473 | 0.1433
0.116 | 0.003895 | 0.1414 | 0.001641 | 0 | 0 | 0.1469 | 0.143
0.117 | 0.003831 | 0.1411 | 0.001595 | 0 | 0 | 0.1465 | 0.1427
0.118 | 0.003769 | 0.1408 | 0.001552 | 0 | 0 | 0.1461 | 0.1424
0.119 | 0.003708 | 0.1405 | 0.001509 | 0 | 0 | 0.1457 | 0.142
0.12 | 0.003648 | 0.1402 | 0.001469 | 0 | 0 | 0.1454 | 0.1417
0.121 | 0.00359 | 0.14 | 0.001429 | 0 | 0 | 0.145 | 0.1414
0.122 | 0.003533 | 0.1397 | 0.001391 | 0 | 0 | 0.1446 | 0.1411
0.123 | 0.003478 | 0.1394 | 0.001355 | 0 | 0 | 0.1442 | 0.1408
0.124 | 0.003424 | 0.1391 | 0.00132 | 0 | 0 | 0.1439 | 0.1405
0.125 | 0.003371 | 0.1389 | 0.001285 | 0 | 0 | 0.1435 | 0.1402
0.126 | 0.003319 | 0.1386 | 0.001252 | 0 | 0 | 0.1432 | 0.1399
0.127 | 0.003268 | 0.1383 | 0.001221 | 0 | 0 | 0.1428 | 0.1395
0.128 | 0.003219 | 0.1381 | 0.00119 | 0 | 0 | 0.1425 | 0.1393
0.129 | 0.003171 | 0.1378 | 0.00116 | 0 | 0 | 0.1421 | 0.139
0.13 | 0.003123 | 0.1375 | 0.001131 | 0 | 0 | 0.1418 | 0.1387
0.131 | 0.003077 | 0.1373 | 0.001103 | 0 | 0 | 0.1414 | 0.1384
0.132 | 0.003032 | 0.137 | 0.001076 | 0 | 0 | 0.1411 | 0.1381
0.133 | 0.002988 | 0.1367 | 0.00105 | 0 | 0 | 0.1408 | 0.1378
0.134 | 0.002945 | 0.1365 | 0.001025 | 0 | 0 | 0.1404 | 0.1375
0.135 | 0.002902 | 0.1362 | 0.001 | 0 | 0 | 0.1401 | 0.1372
0.136 | 0.002861 | 0.136 | 0.0009767 | 0 | 0 | 0.1398 | 0.1369
0.137 | 0.00282 | 0.1357 | 0.0009537 | 0 | 0 | 0.1395 | 0.1367
0.138 | 0.002781 | 0.1354 | 0.0009315 | 0 | 0 | 0.1392 | 0.1364
0.139 | 0.002742 | 0.1352 | 0.0009099 | 0 | 0 | 0.1388 | 0.1361
0.14 | 0.002704 | 0.1349 | 0.0008889 | 0 | 0 | 0.1385 | 0.1358
0.145 | 0.002525 | 0.1337 | 0.0007933 | 0 | 0 | 0.137 | 0.1345
0.15 | 0.002363 | 0.1325 | 0.0007108 | 0 | 0 | 0.1355 | 0.1332
0.155 | 0.002216 | 0.1313 | 0.0006393 | 0 | 0 | 0.1341 | 0.1319
0.16 | 0.002082 | 0.1301 | 0.0005771 | 0 | 0 | 0.1328 | 0.1307
0.165 | 0.00196 | 0.129 | 0.0005226 | 0 | 0 | 0.1315 | 0.1295
0.17 | 0.001848 | 0.1279 | 0.0004748 | 0 | 0 | 0.1302 | 0.1283
0.175 | 0.001746 | 0.1268 | 0.0004326 | 0 | 0 | 0.1289 | 0.1272
0.18 | 0.001652 | 0.1257 | 0.0003953 | 0 | 0 | 0.1278 | 0.1261
0.185 | 0.001565 | 0.1247 | 0.0003621 | 0 | 0 | 0.1266 | 0.125
0.19 | 0.001485 | 0.1237 | 0.0003326 | 0 | 0 | 0.1255 | 0.124
0.195 | 0.001411 | 0.1227 | 0.0003062 | 0 | 0 | 0.1244 | 0.123
0.2 | 0.001342 | 0.1217 | 0.0002825 | 0 | 0 | 0.1233 | 0.122
0.21 | 0.001219 | 0.1198 | 0.0002421 | 0 | 0 | 0.1213 | 0.1201
0.22 | 0.001112 | 0.118 | 0.000209 | 0 | 0 | 0.1194 | 0.1183
0.23 | 0.001018 | 0.1163 | 0.0001818 | 0 | 0 | 0.1175 | 0.1165
0.24 | 0.0009357 | 0.1147 | 0.0001591 | 0 | 0 | 0.1158 | 0.1149
0.25 | 0.000863 | 0.1131 | 0.0001402 | 0 | 0 | 0.1141 | 0.1133
0.26 | 0.0007984 | 0.1116 | 0.0001241 | 0 | 0 | 0.1125 | 0.1117
0.28 | 0.0006892 | 0.1087 | 9.889e-05 | 0 | 0 | 0.1095 | 0.1088
0.3 | 0.0006009 | 0.1061 | 8.02e-05 | 0 | 0 | 0.1068 | 0.1062
0.32 | 0.0005286 | 0.1036 | 6.608e-05 | 0 | 0 | 0.1042 | 0.1037
0.34 | 0.0004685 | 0.1013 | 5.52e-05 | 0 | 0 | 0.1018 | 0.1013
0.36 | 0.0004181 | 0.09911 | 4.668e-05 | 0 | 0 | 0.09958 | 0.09916
0.38 | 0.0003754 | 0.09706 | 3.99e-05 | 0 | 0 | 0.09748 | 0.0971
0.4 | 0.0003389 | 0.09514 | 3.444e-05 | 0 | 0 | 0.09551 | 0.09517
0.42 | 0.0003075 | 0.09331 | 2.999e-05 | 0 | 0 | 0.09365 | 0.09334
0.44 | 0.0002802 | 0.09158 | 2.633e-05 | 0 | 0 | 0.09189 | 0.09161
0.46 | 0.0002564 | 0.08994 | 2.329e-05 | 0 | 0 | 0.09022 | 0.08997
0.48 | 0.0002356 | 0.08838 | 2.074e-05 | 0 | 0 | 0.08864 | 0.08841
0.5 | 0.0002171 | 0.0869 | 1.86e-05 | 0 | 0 | 0.08714 | 0.08692
0.55 | 0.0001795 | 0.08348 | 1.45e-05 | 0 | 0 | 0.08367 | 0.08349
0.6 | 0.0001509 | 0.08041 | 1.159e-05 | 0 | 0 | 0.08057 | 0.08042
0.65 | 0.0001286 | 0.07765 | 9.436e-06 | 0 | 0 | 0.07778 | 0.07766
0.7 | 0.0001109 | 0.07513 | 7.847e-06 | 0 | 0 | 0.07525 | 0.07513
0.75 | 9.665e-05 | 0.07281 | 6.69e-06 | 0 | 0 | 0.07291 | 0.07282
0.8 | 8.494e-05 | 0.07066 | 5.858e-06 | 0 | 0 | 0.07075 | 0.07067
0.85 | 7.523e-05 | 0.06867 | 5.248e-06 | 0 | 0 | 0.06875 | 0.06867
0.9 | 6.709e-05 | 0.06682 | 4.722e-06 | 0 | 0 | 0.06689 | 0.06682
0.95 | 6.022e-05 | 0.06511 | 4.197e-06 | 0 | 0 | 0.06517 | 0.06511
1 | 5.437e-05 | 0.06354 | 3.642e-06 | 0 | 0 | 0.0636 | 0.06354
2 | 1.36e-05 | 0.04408 | 1.049e-06 | 0.0003892 | 0 | 0.04448 | 0.04447
3 | 6.044e-06 | 0.03468 | 5.844e-07 | 0.001113 | 1.213e-05 | 0.03581 | 0.0358
4 | 3.401e-06 | 0.02893 | 4.004e-07 | 0.001805 | 4.953e-05 | 0.03079 | 0.03078
5 | 2.176e-06 | 0.02498 | 3.032e-07 | 0.002421 | 9.871e-05 | 0.0275 | 0.0275
6 | 1.511e-06 | 0.02208 | 2.435e-07 | 0.002976 | 0.0001516 | 0.02521 | 0.02521
7 | 1.11e-06 | 0.01985 | 2.033e-07 | 0.00347 | 0.0002044 | 0.02352 | 0.02352
8 | 8.503e-07 | 0.01806 | 1.743e-07 | 0.003915 | 0.0002557 | 0.02223 | 0.02223
9 | 6.719e-07 | 0.0166 | 1.526e-07 | 0.004319 | 0.0003049 | 0.02123 | 0.02123
10 | 5.442e-07 | 0.01538 | 1.356e-07 | 0.004685 | 0.0003517 | 0.02042 | 0.02042
11 | 4.497e-07 | 0.01435 | 1.22e-07 | 0.00502 | 0.0003962 | 0.01976 | 0.01976
12 | 3.778e-07 | 0.01345 | 1.109e-07 | 0.005325 | 0.0004382 | 0.01922 | 0.01922
13 | 3.22e-07 | 0.01268 | 1.016e-07 | 0.005612 | 0.0004781 | 0.01877 | 0.01877
14 | 2.776e-07 | 0.01199 | 9.375e-08 | 0.005876 | 0.0005159 | 0.01839 | 0.01839
15 | 2.418e-07 | 0.01138 | 8.702e-08 | 0.006126 | 0.0005519 | 0.01806 | 0.01806
16 | 2.126e-07 | 0.01084 | 8.121e-08 | 0.006359 | 0.0005863 | 0.01779 | 0.01779
17 | 1.883e-07 | 0.01035 | 7.611e-08 | 0.006579 | 0.0006191 | 0.01755 | 0.01755
18 | 1.68e-07 | 0.00991 | 7.16e-08 | 0.006785 | 0.0006503 | 0.01735 | 0.01735
19 | 1.507e-07 | 0.009506 | 6.761e-08 | 0.00698 | 0.0006801 | 0.01717 | 0.01717
20 | 1.36e-07 | 0.009137 | 6.404e-08 | 0.007166 | 0.0007087 | 0.01701 | 0.01701

## Lung
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.0739 | 0.167 | 0.341 | 0 | 0 | 0.582 | 0.508
0.021 | 0.0686 | 0.168 | 0.291 | 0 | 0 | 0.528 | 0.459
0.022 | 0.0639 | 0.169 | 0.25 | 0 | 0 | 0.483 | 0.419
0.023 | 0.0596 | 0.17 | 0.216 | 0 | 0 | 0.446 | 0.386
0.024 | 0.0558 | 0.17 | 0.188 | 0 | 0 | 0.414 | 0.359
0.025 | 0.0523 | 0.171 | 0.165 | 0 | 0 | 0.388 | 0.336
0.026 | 0.0491 | 0.171 | 0.145 | 0 | 0 | 0.365 | 0.316
0.027 | 0.0461 | 0.172 | 0.128 | 0 | 0 | 0.346 | 0.3
0.028 | 0.0435 | 0.172 | 0.114 | 0 | 0 | 0.329 | 0.286
0.029 | 0.041 | 0.172 | 0.101 | 0 | 0 | 0.315 | 0.274
0.03 | 0.0387 | 0.173 | 0.0906 | 0 | 0 | 0.302 | 0.263
0.031 | 0.0367 | 0.173 | 0.0814 | 0 | 0 | 0.291 | 0.254
0.032 | 0.0348 | 0.173 | 0.0733 | 0 | 0 | 0.281 | 0.246
0.033 | 0.033 | 0.173 | 0.0662 | 0 | 0 | 0.272 | 0.239
0.034 | 0.0313 | 0.173 | 0.06 | 0 | 0 | 0.264 | 0.233
0.035 | 0.0298 | 0.173 | 0.0546 | 0 | 0 | 0.257 | 0.227
0.036 | 0.0284 | 0.173 | 0.0497 | 0 | 0 | 0.251 | 0.223
0.037 | 0.0271 | 0.173 | 0.0454 | 0 | 0 | 0.245 | 0.218
0.038 | 0.0259 | 0.173 | 0.0416 | 0 | 0 | 0.24 | 0.214
0.039 | 0.0247 | 0.173 | 0.0382 | 0 | 0 | 0.235 | 0.211
0.04 | 0.0236 | 0.172 | 0.0351 | 0 | 0 | 0.231 | 0.208
0.041 | 0.0226 | 0.172 | 0.0324 | 0 | 0 | 0.227 | 0.205
0.042 | 0.0217 | 0.172 | 0.0299 | 0 | 0 | 0.224 | 0.202
0.043 | 0.0208 | 0.172 | 0.0277 | 0 | 0 | 0.22 | 0.199
0.044 | 0.02 | 0.172 | 0.0256 | 0 | 0 | 0.217 | 0.197
0.045 | 0.0192 | 0.171 | 0.0238 | 0 | 0 | 0.214 | 0.195
0.046 | 0.0184 | 0.171 | 0.0221 | 0 | 0 | 0.212 | 0.193
0.047 | 0.0178 | 0.171 | 0.0206 | 0 | 0 | 0.209 | 0.191
0.048 | 0.0171 | 0.171 | 0.0192 | 0 | 0 | 0.207 | 0.19
0.049 | 0.0165 | 0.17 | 0.018 | 0 | 0 | 0.205 | 0.188
0.05 | 0.0159 | 0.17 | 0.0168 | 0 | 0 | 0.203 | 0.187
0.051 | 0.0153 | 0.17 | 0.0157 | 0 | 0 | 0.201 | 0.185
0.052 | 0.0148 | 0.169 | 0.0147 | 0 | 0 | 0.199 | 0.184
0.053 | 0.0143 | 0.169 | 0.0138 | 0 | 0 | 0.197 | 0.183
0.054 | 0.0138 | 0.169 | 0.013 | 0 | 0 | 0.196 | 0.182
0.055 | 0.0134 | 0.169 | 0.0122 | 0 | 0 | 0.194 | 0.181
0.056 | 0.0129 | 0.168 | 0.0115 | 0 | 0 | 0.193 | 0.18
0.057 | 0.0125 | 0.168 | 0.0109 | 0 | 0 | 0.191 | 0.179
0.058 | 0.0121 | 0.168 | 0.0103 | 0 | 0 | 0.19 | 0.178
0.059 | 0.0117 | 0.167 | 0.00969 | 0 | 0 | 0.189 | 0.177
0.06 | 0.0114 | 0.167 | 0.00917 | 0 | 0 | 0.187 | 0.176
0.061 | 0.011 | 0.167 | 0.00868 | 0 | 0 | 0.186 | 0.175
0.062 | 0.0107 | 0.166 | 0.00822 | 0 | 0 | 0.185 | 0.174
0.063 | 0.0104 | 0.166 | 0.0078 | 0 | 0 | 0.184 | 0.174
0.064 | 0.0101 | 0.165 | 0.0074 | 0 | 0 | 0.183 | 0.173
0.065 | 0.00982 | 0.165 | 0.00703 | 0 | 0 | 0.182 | 0.172
0.066 | 0.00955 | 0.165 | 0.00668 | 0 | 0 | 0.181 | 0.171
0.067 | 0.00928 | 0.164 | 0.00636 | 0 | 0 | 0.18 | 0.171
0.068 | 0.00903 | 0.164 | 0.00605 | 0 | 0 | 0.179 | 0.17
0.069 | 0.00879 | 0.164 | 0.00577 | 0 | 0 | 0.178 | 0.17
0.07 | 0.00855 | 0.163 | 0.0055 | 0 | 0 | 0.177 | 0.169
0.071 | 0.00833 | 0.163 | 0.00524 | 0 | 0 | 0.177 | 0.168
0.072 | 0.00811 | 0.163 | 0.00501 | 0 | 0 | 0.176 | 0.168
0.073 | 0.00791 | 0.162 | 0.00478 | 0 | 0 | 0.175 | 0.167
0.074 | 0.00771 | 0.162 | 0.00457 | 0 | 0 | 0.174 | 0.167
0.075 | 0.00752 | 0.162 | 0.00437 | 0 | 0 | 0.174 | 0.166
0.076 | 0.00733 | 0.161 | 0.00418 | 0 | 0 | 0.173 | 0.165
0.077 | 0.00715 | 0.161 | 0.00401 | 0 | 0 | 0.172 | 0.165
0.078 | 0.00698 | 0.161 | 0.00384 | 0 | 0 | 0.171 | 0.164
0.079 | 0.00681 | 0.16 | 0.00368 | 0 | 0 | 0.171 | 0.164
0.08 | 0.00665 | 0.16 | 0.00353 | 0 | 0 | 0.17 | 0.163
0.081 | 0.0065 | 0.16 | 0.00339 | 0 | 0 | 0.169 | 0.163
0.082 | 0.00635 | 0.159 | 0.00325 | 0 | 0 | 0.169 | 0.162
0.083 | 0.00621 | 0.159 | 0.00312 | 0 | 0 | 0.168 | 0.162
0.084 | 0.00607 | 0.159 | 0.003 | 0 | 0 | 0.168 | 0.162
0.085 | 0.00593 | 0.158 | 0.00289 | 0 | 0 | 0.167 | 0.161
0.086 | 0.0058 | 0.158 | 0.00278 | 0 | 0 | 0.166 | 0.161
0.087 | 0.00567 | 0.157 | 0.00267 | 0 | 0 | 0.166 | 0.16
0.088 | 0.00555 | 0.157 | 0.00257 | 0 | 0 | 0.165 | 0.16
0.089 | 0.00543 | 0.157 | 0.00248 | 0 | 0 | 0.165 | 0.159
0.09 | 0.00532 | 0.156 | 0.00239 | 0 | 0 | 0.164 | 0.159
0.091 | 0.00521 | 0.156 | 0.0023 | 0 | 0 | 0.164 | 0.158
0.092 | 0.0051 | 0.156 | 0.00222 | 0 | 0 | 0.163 | 0.158
0.093 | 0.005 | 0.155 | 0.00214 | 0 | 0 | 0.163 | 0.158
0.094 | 0.00489 | 0.155 | 0.00207 | 0 | 0 | 0.162 | 0.157
0.095 | 0.0048 | 0.155 | 0.002 | 0 | 0 | 0.162 | 0.157
0.096 | 0.0047 | 0.154 | 0.00193 | 0 | 0 | 0.161 | 0.156
0.097 | 0.00461 | 0.154 | 0.00186 | 0 | 0 | 0.161 | 0.156
0.098 | 0.00452 | 0.154 | 0.0018 | 0 | 0 | 0.16 | 0.156
0.099 | 0.00443 | 0.153 | 0.00174 | 0 | 0 | 0.16 | 0.155
0.1 | 0.00435 | 0.153 | 0.00169 | 0 | 0 | 0.159 | 0.155
0.101 | 0.00426 | 0.153 | 0.00163 | 0 | 0 | 0.159 | 0.154
0.102 | 0.00418 | 0.152 | 0.00158 | 0 | 0 | 0.158 | 0.154
0.103 | 0.00411 | 0.152 | 0.00153 | 0 | 0 | 0.158 | 0.154
0.104 | 0.00403 | 0.152 | 0.00148 | 0 | 0 | 0.157 | 0.153
0.105 | 0.00396 | 0.152 | 0.00143 | 0 | 0 | 0.157 | 0.153
0.106 | 0.00388 | 0.151 | 0.00139 | 0 | 0 | 0.156 | 0.153
0.107 | 0.00381 | 0.151 | 0.00135 | 0 | 0 | 0.156 | 0.152
0.108 | 0.00375 | 0.151 | 0.00131 | 0 | 0 | 0.156 | 0.152
0.109 | 0.00368 | 0.15 | 0.00127 | 0 | 0 | 0.155 | 0.152
0.11 | 0.00362 | 0.15 | 0.00123 | 0 | 0 | 0.155 | 0.151
0.111 | 0.00355 | 0.15 | 0.00119 | 0 | 0 | 0.154 | 0.151
0.112 | 0.00349 | 0.149 | 0.00116 | 0 | 0 | 0.154 | 0.15
0.113 | 0.00343 | 0.149 | 0.00113 | 0 | 0 | 0.154 | 0.15
0.114 | 0.00337 | 0.149 | 0.00109 | 0 | 0 | 0.153 | 0.15
0.115 | 0.00332 | 0.148 | 0.00106 | 0 | 0 | 0.153 | 0.149
0.116 | 0.00326 | 0.148 | 0.00103 | 0 | 0 | 0.152 | 0.149
0.117 | 0.00321 | 0.148 | 0.001 | 0 | 0 | 0.152 | 0.149
0.118 | 0.00316 | 0.148 | 0.000976 | 0 | 0 | 0.152 | 0.148
0.119 | 0.00311 | 0.147 | 0.000949 | 0 | 0 | 0.151 | 0.148
0.12 | 0.00306 | 0.147 | 0.000923 | 0 | 0 | 0.151 | 0.148
0.121 | 0.00301 | 0.147 | 0.000898 | 0 | 0 | 0.151 | 0.148
0.122 | 0.00296 | 0.146 | 0.000874 | 0 | 0 | 0.15 | 0.147
0.123 | 0.00291 | 0.146 | 0.000851 | 0 | 0 | 0.15 | 0.147
0.124 | 0.00287 | 0.146 | 0.000829 | 0 | 0 | 0.149 | 0.147
0.125 | 0.00282 | 0.145 | 0.000807 | 0 | 0 | 0.149 | 0.146
0.126 | 0.00278 | 0.145 | 0.000786 | 0 | 0 | 0.149 | 0.146
0.127 | 0.00274 | 0.145 | 0.000766 | 0 | 0 | 0.148 | 0.146
0.128 | 0.00269 | 0.145 | 0.000747 | 0 | 0 | 0.148 | 0.145
0.129 | 0.00265 | 0.144 | 0.000728 | 0 | 0 | 0.148 | 0.145
0.13 | 0.00261 | 0.144 | 0.00071 | 0 | 0 | 0.147 | 0.145
0.131 | 0.00258 | 0.144 | 0.000692 | 0 | 0 | 0.147 | 0.144
0.132 | 0.00254 | 0.143 | 0.000675 | 0 | 0 | 0.147 | 0.144
0.133 | 0.0025 | 0.143 | 0.000658 | 0 | 0 | 0.146 | 0.144
0.134 | 0.00246 | 0.143 | 0.000642 | 0 | 0 | 0.146 | 0.144
0.135 | 0.00243 | 0.143 | 0.000627 | 0 | 0 | 0.146 | 0.143
0.136 | 0.00239 | 0.142 | 0.000612 | 0 | 0 | 0.145 | 0.143
0.137 | 0.00236 | 0.142 | 0.000597 | 0 | 0 | 0.145 | 0.143
0.138 | 0.00233 | 0.142 | 0.000583 | 0 | 0 | 0.145 | 0.142
0.139 | 0.00229 | 0.142 | 0.00057 | 0 | 0 | 0.144 | 0.142
0.14 | 0.00226 | 0.141 | 0.000557 | 0 | 0 | 0.144 | 0.142
0.145 | 0.00211 | 0.14 | 0.000496 | 0 | 0 | 0.143 | 0.14
0.15 | 0.00198 | 0.139 | 0.000444 | 0 | 0 | 0.141 | 0.139
0.155 | 0.00185 | 0.137 | 0.000399 | 0 | 0 | 0.14 | 0.138
0.16 | 0.00174 | 0.136 | 0.00036 | 0 | 0 | 0.138 | 0.137
0.165 | 0.00164 | 0.135 | 0.000326 | 0 | 0 | 0.137 | 0.135
0.17 | 0.00155 | 0.134 | 0.000296 | 0 | 0 | 0.136 | 0.134
0.175 | 0.00146 | 0.133 | 0.000269 | 0 | 0 | 0.134 | 0.133
0.18 | 0.00138 | 0.132 | 0.000246 | 0 | 0 | 0.133 | 0.132
0.185 | 0.00131 | 0.13 | 0.000225 | 0 | 0 | 0.132 | 0.131
0.19 | 0.00124 | 0.129 | 0.000207 | 0 | 0 | 0.131 | 0.13
0.195 | 0.00118 | 0.128 | 0.00019 | 0 | 0 | 0.13 | 0.129
0.2 | 0.00112 | 0.127 | 0.000175 | 0 | 0 | 0.129 | 0.128
0.21 | 0.00102 | 0.125 | 0.00015 | 0 | 0 | 0.127 | 0.126
0.22 | 0.000928 | 0.124 | 0.000129 | 0 | 0 | 0.125 | 0.124
0.23 | 0.00085 | 0.122 | 0.000112 | 0 | 0 | 0.123 | 0.122
0.24 | 0.000781 | 0.12 | 9.83e-05 | 0 | 0 | 0.121 | 0.12
0.25 | 0.00072 | 0.118 | 8.65e-05 | 0 | 0 | 0.119 | 0.118
0.26 | 0.000666 | 0.117 | 7.66e-05 | 0 | 0 | 0.118 | 0.117
0.28 | 0.000575 | 0.114 | 6.09e-05 | 0 | 0 | 0.114 | 0.114
0.3 | 0.000501 | 0.111 | 4.93e-05 | 0 | 0 | 0.112 | 0.111
0.32 | 0.000441 | 0.108 | 4.06e-05 | 0 | 0 | 0.109 | 0.108
0.34 | 0.000391 | 0.106 | 3.39e-05 | 0 | 0 | 0.106 | 0.106
0.36 | 0.000349 | 0.104 | 2.86e-05 | 0 | 0 | 0.104 | 0.104
0.38 | 0.000313 | 0.102 | 2.45e-05 | 0 | 0 | 0.102 | 0.102
0.4 | 0.000283 | 0.0995 | 2.11e-05 | 0 | 0 | 0.0998 | 0.0995
0.42 | 0.000256 | 0.0976 | 1.84e-05 | 0 | 0 | 0.0979 | 0.0976
0.44 | 0.000234 | 0.0958 | 1.61e-05 | 0 | 0 | 0.096 | 0.0958
0.46 | 0.000214 | 0.0941 | 1.42e-05 | 0 | 0 | 0.0943 | 0.0941
0.48 | 0.000197 | 0.0924 | 1.27e-05 | 0 | 0 | 0.0926 | 0.0925
0.5 | 0.000181 | 0.0909 | 1.14e-05 | 0 | 0 | 0.0911 | 0.0909
0.55 | 0.00015 | 0.0873 | 8.86e-06 | 0 | 0 | 0.0875 | 0.0873
0.6 | 0.000126 | 0.0841 | 7.07e-06 | 0 | 0 | 0.0842 | 0.0841
0.65 | 0.000107 | 0.0812 | 5.74e-06 | 0 | 0 | 0.0813 | 0.0812
0.7 | 9.25e-05 | 0.0785 | 4.77e-06 | 0 | 0 | 0.0786 | 0.0785
0.75 | 8.06e-05 | 0.0761 | 4.06e-06 | 0 | 0 | 0.0762 | 0.0761
0.8 | 7.08e-05 | 0.0739 | 3.57e-06 | 0 | 0 | 0.074 | 0.0739
0.85 | 6.27e-05 | 0.0718 | 3.21e-06 | 0 | 0 | 0.0719 | 0.0718
0.9 | 5.6e-05 | 0.0699 | 2.9e-06 | 0 | 0 | 0.0699 | 0.0699
0.95 | 5.02e-05 | 0.0681 | 2.58e-06 | 0 | 0 | 0.0681 | 0.0681
1 | 4.53e-05 | 0.0664 | 2.22e-06 | 0 | 0 | 0.0665 | 0.0664
2 | 1.13e-05 | 0.0461 | 6.39e-07 | 0.000347 | 0 | 0.0464 | 0.0464
3 | 5.04e-06 | 0.0363 | 3.58e-07 | 0.000992 | 1.27e-05 | 0.0373 | 0.0373
4 | 2.84e-06 | 0.0302 | 2.46e-07 | 0.00161 | 5.18e-05 | 0.0319 | 0.0319
5 | 1.81e-06 | 0.0261 | 1.86e-07 | 0.00216 | 0.000103 | 0.0284 | 0.0284
6 | 1.26e-06 | 0.0231 | 1.5e-07 | 0.00266 | 0.000159 | 0.0259 | 0.0259
7 | 9.26e-07 | 0.0208 | 1.25e-07 | 0.0031 | 0.000214 | 0.0241 | 0.0241
8 | 7.09e-07 | 0.0189 | 1.08e-07 | 0.0035 | 0.000267 | 0.0226 | 0.0226
9 | 5.6e-07 | 0.0174 | 9.42e-08 | 0.00386 | 0.000319 | 0.0215 | 0.0215
10 | 4.54e-07 | 0.0161 | 8.38e-08 | 0.00418 | 0.000368 | 0.0206 | 0.0206
11 | 3.75e-07 | 0.015 | 7.54e-08 | 0.00448 | 0.000414 | 0.0199 | 0.0199
12 | 3.15e-07 | 0.0141 | 6.86e-08 | 0.00476 | 0.000458 | 0.0193 | 0.0193
13 | 2.68e-07 | 0.0133 | 6.28e-08 | 0.00501 | 0.0005 | 0.0188 | 0.0188
14 | 2.31e-07 | 0.0125 | 5.8e-08 | 0.00525 | 0.00054 | 0.0183 | 0.0183
15 | 2.02e-07 | 0.0119 | 5.39e-08 | 0.00547 | 0.000577 | 0.018 | 0.018
16 | 1.77e-07 | 0.0113 | 5.03e-08 | 0.00568 | 0.000613 | 0.0176 | 0.0176
17 | 1.57e-07 | 0.0108 | 4.71e-08 | 0.00588 | 0.000647 | 0.0174 | 0.0174
18 | 1.4e-07 | 0.0104 | 4.43e-08 | 0.00606 | 0.00068 | 0.0171 | 0.0171
19 | 1.26e-07 | 0.00994 | 4.19e-08 | 0.00624 | 0.000711 | 0.0169 | 0.0169
20 | 1.13e-07 | 0.00955 | 3.97e-08 | 0.0064 | 0.000741 | 0.0167 | 0.0167

## PMP
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.05643 | 0.1887 | 0.1864 | 0 | 0 | 0.4316 | 0.3751
0.021 | 0.05235 | 0.1896 | 0.1588 | 0 | 0 | 0.4007 | 0.3483
0.022 | 0.04867 | 0.1903 | 0.1362 | 0 | 0 | 0.3752 | 0.3265
0.023 | 0.04535 | 0.1909 | 0.1176 | 0 | 0 | 0.3539 | 0.3085
0.024 | 0.04235 | 0.1914 | 0.1022 | 0 | 0 | 0.336 | 0.2936
0.025 | 0.03964 | 0.1917 | 0.08934 | 0 | 0 | 0.3207 | 0.2811
0.026 | 0.03716 | 0.192 | 0.07848 | 0 | 0 | 0.3077 | 0.2705
0.027 | 0.03491 | 0.1923 | 0.06927 | 0 | 0 | 0.2964 | 0.2615
0.028 | 0.03285 | 0.1924 | 0.06142 | 0 | 0 | 0.2867 | 0.2538
0.029 | 0.03097 | 0.1925 | 0.05468 | 0 | 0 | 0.2782 | 0.2472
0.03 | 0.02925 | 0.1925 | 0.04887 | 0 | 0 | 0.2707 | 0.2414
0.031 | 0.02766 | 0.1925 | 0.04383 | 0 | 0 | 0.264 | 0.2364
0.032 | 0.0262 | 0.1925 | 0.03945 | 0 | 0 | 0.2581 | 0.2319
0.033 | 0.02485 | 0.1924 | 0.03562 | 0 | 0 | 0.2529 | 0.228
0.034 | 0.0236 | 0.1923 | 0.03225 | 0 | 0 | 0.2481 | 0.2245
0.035 | 0.02245 | 0.1921 | 0.02929 | 0 | 0 | 0.2438 | 0.2214
0.036 | 0.02137 | 0.1919 | 0.02667 | 0 | 0 | 0.24 | 0.2186
0.037 | 0.02037 | 0.1917 | 0.02435 | 0 | 0 | 0.2364 | 0.2161
0.038 | 0.01944 | 0.1915 | 0.02228 | 0 | 0 | 0.2332 | 0.2138
0.039 | 0.01857 | 0.1912 | 0.02043 | 0 | 0 | 0.2302 | 0.2117
0.04 | 0.01776 | 0.191 | 0.01878 | 0 | 0 | 0.2275 | 0.2097
0.041 | 0.017 | 0.1907 | 0.0173 | 0 | 0 | 0.225 | 0.208
0.042 | 0.01628 | 0.1904 | 0.01597 | 0 | 0 | 0.2226 | 0.2063
0.043 | 0.01562 | 0.1901 | 0.01476 | 0 | 0 | 0.2204 | 0.2048
0.044 | 0.01499 | 0.1897 | 0.01367 | 0 | 0 | 0.2184 | 0.2034
0.045 | 0.01439 | 0.1894 | 0.01269 | 0 | 0 | 0.2165 | 0.2021
0.046 | 0.01384 | 0.1891 | 0.01179 | 0 | 0 | 0.2147 | 0.2008
0.047 | 0.01331 | 0.1887 | 0.01097 | 0 | 0 | 0.213 | 0.1997
0.048 | 0.01281 | 0.1883 | 0.01023 | 0 | 0 | 0.2114 | 0.1986
0.049 | 0.01234 | 0.188 | 0.00955 | 0 | 0 | 0.2099 | 0.1975
0.05 | 0.0119 | 0.1876 | 0.008927 | 0 | 0 | 0.2084 | 0.1965
0.051 | 0.01147 | 0.1872 | 0.008356 | 0 | 0 | 0.207 | 0.1956
0.052 | 0.01107 | 0.1868 | 0.007832 | 0 | 0 | 0.2057 | 0.1947
0.053 | 0.01069 | 0.1864 | 0.007349 | 0 | 0 | 0.2045 | 0.1938
0.054 | 0.01033 | 0.186 | 0.006905 | 0 | 0 | 0.2033 | 0.1929
0.055 | 0.009992 | 0.1856 | 0.006494 | 0 | 0 | 0.2021 | 0.1921
0.056 | 0.009666 | 0.1852 | 0.006115 | 0 | 0 | 0.201 | 0.1914
0.057 | 0.009355 | 0.1848 | 0.005764 | 0 | 0 | 0.2 | 0.1906
0.058 | 0.009059 | 0.1844 | 0.005439 | 0 | 0 | 0.1989 | 0.1899
0.059 | 0.008777 | 0.184 | 0.005137 | 0 | 0 | 0.1979 | 0.1892
0.06 | 0.008508 | 0.1836 | 0.004857 | 0 | 0 | 0.197 | 0.1885
0.061 | 0.00825 | 0.1832 | 0.004596 | 0 | 0 | 0.1961 | 0.1878
0.062 | 0.008005 | 0.1828 | 0.004353 | 0 | 0 | 0.1952 | 0.1871
0.063 | 0.007769 | 0.1824 | 0.004126 | 0 | 0 | 0.1943 | 0.1865
0.064 | 0.007544 | 0.182 | 0.003915 | 0 | 0 | 0.1934 | 0.1859
0.065 | 0.007329 | 0.1816 | 0.003717 | 0 | 0 | 0.1926 | 0.1853
0.066 | 0.007123 | 0.1812 | 0.003532 | 0 | 0 | 0.1918 | 0.1847
0.067 | 0.006925 | 0.1807 | 0.003359 | 0 | 0 | 0.191 | 0.1841
0.068 | 0.006735 | 0.1803 | 0.003197 | 0 | 0 | 0.1903 | 0.1835
0.069 | 0.006553 | 0.1799 | 0.003045 | 0 | 0 | 0.1895 | 0.183
0.07 | 0.006377 | 0.1795 | 0.002902 | 0 | 0 | 0.1888 | 0.1824
0.071 | 0.006209 | 0.1791 | 0.002768 | 0 | 0 | 0.1881 | 0.1819
0.072 | 0.006048 | 0.1787 | 0.002641 | 0 | 0 | 0.1874 | 0.1813
0.073 | 0.005892 | 0.1783 | 0.002522 | 0 | 0 | 0.1867 | 0.1808
0.074 | 0.005742 | 0.1779 | 0.00241 | 0 | 0 | 0.186 | 0.1803
0.075 | 0.005598 | 0.1775 | 0.002305 | 0 | 0 | 0.1854 | 0.1798
0.076 | 0.005459 | 0.1771 | 0.002205 | 0 | 0 | 0.1847 | 0.1793
0.077 | 0.005326 | 0.1767 | 0.002111 | 0 | 0 | 0.1841 | 0.1788
0.078 | 0.005197 | 0.1763 | 0.002022 | 0 | 0 | 0.1835 | 0.1783
0.079 | 0.005072 | 0.1759 | 0.001938 | 0 | 0 | 0.1829 | 0.1778
0.08 | 0.004952 | 0.1755 | 0.001858 | 0 | 0 | 0.1823 | 0.1773
0.081 | 0.004836 | 0.1751 | 0.001783 | 0 | 0 | 0.1817 | 0.1769
0.082 | 0.004724 | 0.1747 | 0.001711 | 0 | 0 | 0.1811 | 0.1764
0.083 | 0.004616 | 0.1743 | 0.001643 | 0 | 0 | 0.1805 | 0.1759
0.084 | 0.004511 | 0.1739 | 0.001579 | 0 | 0 | 0.18 | 0.1755
0.085 | 0.00441 | 0.1735 | 0.001518 | 0 | 0 | 0.1794 | 0.175
0.086 | 0.004312 | 0.1731 | 0.00146 | 0 | 0 | 0.1789 | 0.1746
0.087 | 0.004218 | 0.1727 | 0.001404 | 0 | 0 | 0.1783 | 0.1741
0.088 | 0.004126 | 0.1723 | 0.001352 | 0 | 0 | 0.1778 | 0.1737
0.089 | 0.004038 | 0.1719 | 0.001302 | 0 | 0 | 0.1773 | 0.1732
0.09 | 0.003952 | 0.1716 | 0.001254 | 0 | 0 | 0.1768 | 0.1728
0.091 | 0.003869 | 0.1712 | 0.001209 | 0 | 0 | 0.1763 | 0.1724
0.092 | 0.003788 | 0.1708 | 0.001166 | 0 | 0 | 0.1757 | 0.172
0.093 | 0.00371 | 0.1704 | 0.001125 | 0 | 0 | 0.1752 | 0.1715
0.094 | 0.003634 | 0.17 | 0.001085 | 0 | 0 | 0.1748 | 0.1711
0.095 | 0.003561 | 0.1697 | 0.001048 | 0 | 0 | 0.1743 | 0.1707
0.096 | 0.003489 | 0.1693 | 0.001012 | 0 | 0 | 0.1738 | 0.1703
0.097 | 0.00342 | 0.1689 | 0.0009774 | 0 | 0 | 0.1733 | 0.1699
0.098 | 0.003353 | 0.1685 | 0.0009446 | 0 | 0 | 0.1728 | 0.1695
0.099 | 0.003288 | 0.1682 | 0.0009133 | 0 | 0 | 0.1724 | 0.1691
0.1 | 0.003224 | 0.1678 | 0.0008833 | 0 | 0 | 0.1719 | 0.1687
0.101 | 0.003163 | 0.1674 | 0.0008545 | 0 | 0 | 0.1715 | 0.1683
0.102 | 0.003103 | 0.1671 | 0.000827 | 0 | 0 | 0.171 | 0.1679
0.103 | 0.003045 | 0.1667 | 0.0008007 | 0 | 0 | 0.1706 | 0.1675
0.104 | 0.002988 | 0.1664 | 0.0007754 | 0 | 0 | 0.1701 | 0.1671
0.105 | 0.002933 | 0.166 | 0.0007511 | 0 | 0 | 0.1697 | 0.1668
0.106 | 0.00288 | 0.1656 | 0.0007279 | 0 | 0 | 0.1693 | 0.1664
0.107 | 0.002828 | 0.1653 | 0.0007056 | 0 | 0 | 0.1688 | 0.166
0.108 | 0.002777 | 0.1649 | 0.0006841 | 0 | 0 | 0.1684 | 0.1656
0.109 | 0.002728 | 0.1646 | 0.0006635 | 0 | 0 | 0.168 | 0.1653
0.11 | 0.00268 | 0.1642 | 0.0006438 | 0 | 0 | 0.1676 | 0.1649
0.111 | 0.002633 | 0.1639 | 0.0006247 | 0 | 0 | 0.1672 | 0.1645
0.112 | 0.002588 | 0.1636 | 0.0006065 | 0 | 0 | 0.1667 | 0.1642
0.113 | 0.002543 | 0.1632 | 0.0005889 | 0 | 0 | 0.1663 | 0.1638
0.114 | 0.0025 | 0.1629 | 0.000572 | 0 | 0 | 0.1659 | 0.1634
0.115 | 0.002458 | 0.1625 | 0.0005557 | 0 | 0 | 0.1655 | 0.1631
0.116 | 0.002417 | 0.1622 | 0.00054 | 0 | 0 | 0.1651 | 0.1627
0.117 | 0.002376 | 0.1619 | 0.0005249 | 0 | 0 | 0.1648 | 0.1624
0.118 | 0.002337 | 0.1615 | 0.0005103 | 0 | 0 | 0.1644 | 0.162
0.119 | 0.002299 | 0.1612 | 0.0004963 | 0 | 0 | 0.164 | 0.1617
0.12 | 0.002262 | 0.1609 | 0.0004827 | 0 | 0 | 0.1636 | 0.1613
0.121 | 0.002226 | 0.1605 | 0.0004697 | 0 | 0 | 0.1632 | 0.161
0.122 | 0.00219 | 0.1602 | 0.0004571 | 0 | 0 | 0.1629 | 0.1607
0.123 | 0.002155 | 0.1599 | 0.0004449 | 0 | 0 | 0.1625 | 0.1603
0.124 | 0.002121 | 0.1596 | 0.0004332 | 0 | 0 | 0.1621 | 0.16
0.125 | 0.002088 | 0.1592 | 0.0004219 | 0 | 0 | 0.1617 | 0.1597
0.126 | 0.002056 | 0.1589 | 0.0004109 | 0 | 0 | 0.1614 | 0.1593
0.127 | 0.002025 | 0.1586 | 0.0004004 | 0 | 0 | 0.161 | 0.159
0.128 | 0.001994 | 0.1583 | 0.0003902 | 0 | 0 | 0.1607 | 0.1587
0.129 | 0.001964 | 0.158 | 0.0003803 | 0 | 0 | 0.1603 | 0.1583
0.13 | 0.001934 | 0.1577 | 0.0003707 | 0 | 0 | 0.16 | 0.158
0.131 | 0.001905 | 0.1573 | 0.0003615 | 0 | 0 | 0.1596 | 0.1577
0.132 | 0.001877 | 0.157 | 0.0003526 | 0 | 0 | 0.1593 | 0.1574
0.133 | 0.00185 | 0.1567 | 0.0003439 | 0 | 0 | 0.1589 | 0.1571
0.134 | 0.001823 | 0.1564 | 0.0003355 | 0 | 0 | 0.1586 | 0.1568
0.135 | 0.001796 | 0.1561 | 0.0003274 | 0 | 0 | 0.1582 | 0.1564
0.136 | 0.00177 | 0.1558 | 0.0003196 | 0 | 0 | 0.1579 | 0.1561
0.137 | 0.001745 | 0.1555 | 0.000312 | 0 | 0 | 0.1576 | 0.1558
0.138 | 0.00172 | 0.1552 | 0.0003046 | 0 | 0 | 0.1572 | 0.1555
0.139 | 0.001696 | 0.1549 | 0.0002975 | 0 | 0 | 0.1569 | 0.1552
0.14 | 0.001672 | 0.1546 | 0.0002906 | 0 | 0 | 0.1566 | 0.1549
0.145 | 0.001561 | 0.1531 | 0.000259 | 0 | 0 | 0.155 | 0.1534
0.15 | 0.00146 | 0.1517 | 0.0002317 | 0 | 0 | 0.1534 | 0.152
0.155 | 0.001369 | 0.1503 | 0.0002082 | 0 | 0 | 0.1519 | 0.1505
0.16 | 0.001286 | 0.149 | 0.0001877 | 0 | 0 | 0.1505 | 0.1492
0.165 | 0.00121 | 0.1477 | 0.0001698 | 0 | 0 | 0.149 | 0.1478
0.17 | 0.001141 | 0.1464 | 0.000154 | 0 | 0 | 0.1477 | 0.1465
0.175 | 0.001077 | 0.1451 | 0.0001402 | 0 | 0 | 0.1463 | 0.1453
0.18 | 0.001019 | 0.1439 | 0.0001279 | 0 | 0 | 0.145 | 0.144
0.185 | 0.000965 | 0.1427 | 0.0001171 | 0 | 0 | 0.1438 | 0.1428
0.19 | 0.0009154 | 0.1415 | 0.0001074 | 0 | 0 | 0.1425 | 0.1416
0.195 | 0.0008695 | 0.1404 | 9.88e-05 | 0 | 0 | 0.1413 | 0.1405
0.2 | 0.0008269 | 0.1393 | 9.107e-05 | 0 | 0 | 0.1402 | 0.1394
0.21 | 0.0007507 | 0.1371 | 7.789e-05 | 0 | 0 | 0.1379 | 0.1372
0.22 | 0.0006845 | 0.1351 | 6.714e-05 | 0 | 0 | 0.1358 | 0.1351
0.23 | 0.0006267 | 0.1331 | 5.831e-05 | 0 | 0 | 0.1338 | 0.1331
0.24 | 0.0005759 | 0.1312 | 5.097e-05 | 0 | 0 | 0.1318 | 0.1312
0.25 | 0.0005311 | 0.1294 | 4.484e-05 | 0 | 0 | 0.13 | 0.1294
0.26 | 0.0004912 | 0.1276 | 3.966e-05 | 0 | 0 | 0.1282 | 0.1277
0.28 | 0.0004239 | 0.1243 | 3.153e-05 | 0 | 0 | 0.1248 | 0.1244
0.3 | 0.0003695 | 0.1213 | 2.552e-05 | 0 | 0 | 0.1217 | 0.1213
0.32 | 0.0003249 | 0.1184 | 2.1e-05 | 0 | 0 | 0.1188 | 0.1185
0.34 | 0.0002879 | 0.1158 | 1.751e-05 | 0 | 0 | 0.1161 | 0.1158
0.36 | 0.0002569 | 0.1133 | 1.479e-05 | 0 | 0 | 0.1136 | 0.1133
0.38 | 0.0002307 | 0.1109 | 1.263e-05 | 0 | 0 | 0.1112 | 0.111
0.4 | 0.0002082 | 0.1087 | 1.089e-05 | 0 | 0 | 0.109 | 0.1087
0.42 | 0.0001889 | 0.1066 | 9.472e-06 | 0 | 0 | 0.1068 | 0.1067
0.44 | 0.0001722 | 0.1047 | 8.307e-06 | 0 | 0 | 0.1049 | 0.1047
0.46 | 0.0001575 | 0.1028 | 7.341e-06 | 0 | 0 | 0.103 | 0.1028
0.48 | 0.0001447 | 0.101 | 6.535e-06 | 0 | 0 | 0.1012 | 0.101
0.5 | 0.0001334 | 0.09933 | 5.857e-06 | 0 | 0 | 0.09947 | 0.09934
0.55 | 0.0001103 | 0.0954 | 4.566e-06 | 0 | 0 | 0.09552 | 0.09541
0.6 | 9.268e-05 | 0.09188 | 3.642e-06 | 0 | 0 | 0.09198 | 0.09189
0.65 | 7.898e-05 | 0.08871 | 2.95e-06 | 0 | 0 | 0.08879 | 0.08871
0.7 | 6.81e-05 | 0.08582 | 2.444e-06 | 0 | 0 | 0.0859 | 0.08583
0.75 | 5.932e-05 | 0.08317 | 2.085e-06 | 0 | 0 | 0.08324 | 0.08318
0.8 | 5.214e-05 | 0.08072 | 1.837e-06 | 0 | 0 | 0.08078 | 0.08072
0.85 | 4.619e-05 | 0.07844 | 1.664e-06 | 0 | 0 | 0.07849 | 0.07845
0.9 | 4.12e-05 | 0.07634 | 1.509e-06 | 0 | 0 | 0.07638 | 0.07634
0.95 | 3.698e-05 | 0.07439 | 1.338e-06 | 0 | 0 | 0.07442 | 0.07439
1 | 3.338e-05 | 0.07259 | 1.142e-06 | 0 | 0 | 0.07263 | 0.07259
2 | 8.347e-06 | 0.05035 | 3.277e-07 | 0.000288 | 0 | 0.05064 | 0.05064
3 | 3.71e-06 | 0.03961 | 1.839e-07 | 0.0008247 | 1.386e-05 | 0.04045 | 0.04045
4 | 2.087e-06 | 0.03304 | 1.265e-07 | 0.001339 | 5.658e-05 | 0.03444 | 0.03444
5 | 1.336e-06 | 0.02854 | 9.614e-08 | 0.001797 | 0.0001128 | 0.03045 | 0.03045
6 | 9.277e-07 | 0.02522 | 7.738e-08 | 0.00221 | 0.0001732 | 0.02761 | 0.0276
7 | 6.813e-07 | 0.02267 | 6.471e-08 | 0.002577 | 0.0002335 | 0.02548 | 0.02548
8 | 5.218e-07 | 0.02063 | 5.556e-08 | 0.002909 | 0.0002921 | 0.02383 | 0.02383
9 | 4.122e-07 | 0.01897 | 4.869e-08 | 0.00321 | 0.0003482 | 0.02252 | 0.02252
10 | 3.339e-07 | 0.01757 | 4.333e-08 | 0.003483 | 0.0004018 | 0.02145 | 0.02145
11 | 2.76e-07 | 0.01639 | 3.899e-08 | 0.003732 | 0.0004526 | 0.02057 | 0.02057
12 | 2.319e-07 | 0.01537 | 3.546e-08 | 0.003961 | 0.0005006 | 0.01983 | 0.01983
13 | 1.976e-07 | 0.01448 | 3.251e-08 | 0.004174 | 0.000546 | 0.0192 | 0.0192
14 | 1.704e-07 | 0.0137 | 3.001e-08 | 0.004372 | 0.0005896 | 0.01866 | 0.01866
15 | 1.484e-07 | 0.01301 | 2.787e-08 | 0.004555 | 0.0006308 | 0.01819 | 0.01819
16 | 1.304e-07 | 0.01238 | 2.601e-08 | 0.004732 | 0.0006698 | 0.01778 | 0.01778
17 | 1.155e-07 | 0.01182 | 2.439e-08 | 0.004897 | 0.0007073 | 0.01743 | 0.01743
18 | 1.031e-07 | 0.01132 | 2.295e-08 | 0.00505 | 0.0007431 | 0.01711 | 0.01711
19 | 9.249e-08 | 0.01086 | 2.168e-08 | 0.005196 | 0.0007773 | 0.01683 | 0.01683
20 | 8.347e-08 | 0.01044 | 2.053e-08 | 0.005336 | 0.0008101 | 0.01658 | 0.01658

## LDPE
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.05643 | 0.1887 | 0.1864 | 0 | 0 | 0.4316 | 0.3751
0.021 | 0.05235 | 0.1896 | 0.1588 | 0 | 0 | 0.4007 | 0.3483
0.022 | 0.04867 | 0.1903 | 0.1362 | 0 | 0 | 0.3752 | 0.3265
0.023 | 0.04535 | 0.1909 | 0.1176 | 0 | 0 | 0.3539 | 0.3085
0.024 | 0.04235 | 0.1914 | 0.1022 | 0 | 0 | 0.336 | 0.2936
0.025 | 0.03964 | 0.1917 | 0.08934 | 0 | 0 | 0.3207 | 0.2811
0.026 | 0.03716 | 0.192 | 0.07848 | 0 | 0 | 0.3077 | 0.2705
0.027 | 0.03491 | 0.1923 | 0.06927 | 0 | 0 | 0.2964 | 0.2615
0.028 | 0.03285 | 0.1924 | 0.06142 | 0 | 0 | 0.2867 | 0.2538
0.029 | 0.03097 | 0.1925 | 0.05468 | 0 | 0 | 0.2782 | 0.2472
0.03 | 0.02925 | 0.1925 | 0.04887 | 0 | 0 | 0.2707 | 0.2414
0.031 | 0.02766 | 0.1925 | 0.04383 | 0 | 0 | 0.264 | 0.2364
0.032 | 0.0262 | 0.1925 | 0.03945 | 0 | 0 | 0.2581 | 0.2319
0.033 | 0.02485 | 0.1924 | 0.03562 | 0 | 0 | 0.2529 | 0.228
0.034 | 0.0236 | 0.1923 | 0.03225 | 0 | 0 | 0.2481 | 0.2245
0.035 | 0.02245 | 0.1921 | 0.02929 | 0 | 0 | 0.2438 | 0.2214
0.036 | 0.02137 | 0.1919 | 0.02667 | 0 | 0 | 0.24 | 0.2186
0.037 | 0.02037 | 0.1917 | 0.02435 | 0 | 0 | 0.2364 | 0.2161
0.038 | 0.01944 | 0.1915 | 0.02228 | 0 | 0 | 0.2332 | 0.2138
0.039 | 0.01857 | 0.1912 | 0.02043 | 0 | 0 | 0.2302 | 0.2117
0.04 | 0.01776 | 0.191 | 0.01878 | 0 | 0 | 0.2275 | 0.2097
0.041 | 0.017 | 0.1907 | 0.0173 | 0 | 0 | 0.225 | 0.208
0.042 | 0.01628 | 0.1904 | 0.01597 | 0 | 0 | 0.2226 | 0.2063
0.043 | 0.01562 | 0.1901 | 0.01476 | 0 | 0 | 0.2204 | 0.2048
0.044 | 0.01499 | 0.1897 | 0.01367 | 0 | 0 | 0.2184 | 0.2034
0.045 | 0.01439 | 0.1894 | 0.01269 | 0 | 0 | 0.2165 | 0.2021
0.046 | 0.01384 | 0.1891 | 0.01179 | 0 | 0 | 0.2147 | 0.2008
0.047 | 0.01331 | 0.1887 | 0.01097 | 0 | 0 | 0.213 | 0.1997
0.048 | 0.01281 | 0.1883 | 0.01023 | 0 | 0 | 0.2114 | 0.1986
0.049 | 0.01234 | 0.188 | 0.00955 | 0 | 0 | 0.2099 | 0.1975
0.05 | 0.0119 | 0.1876 | 0.008927 | 0 | 0 | 0.2084 | 0.1965
0.051 | 0.01147 | 0.1872 | 0.008356 | 0 | 0 | 0.207 | 0.1956
0.052 | 0.01107 | 0.1868 | 0.007832 | 0 | 0 | 0.2057 | 0.1947
0.053 | 0.01069 | 0.1864 | 0.007349 | 0 | 0 | 0.2045 | 0.1938
0.054 | 0.01033 | 0.186 | 0.006905 | 0 | 0 | 0.2033 | 0.1929
0.055 | 0.009992 | 0.1856 | 0.006494 | 0 | 0 | 0.2021 | 0.1921
0.056 | 0.009666 | 0.1852 | 0.006115 | 0 | 0 | 0.201 | 0.1914
0.057 | 0.009355 | 0.1848 | 0.005764 | 0 | 0 | 0.2 | 0.1906
0.058 | 0.009059 | 0.1844 | 0.005439 | 0 | 0 | 0.1989 | 0.1899
0.059 | 0.008777 | 0.184 | 0.005137 | 0 | 0 | 0.1979 | 0.1892
0.06 | 0.008508 | 0.1836 | 0.004857 | 0 | 0 | 0.197 | 0.1885
0.061 | 0.00825 | 0.1832 | 0.004596 | 0 | 0 | 0.1961 | 0.1878
0.062 | 0.008005 | 0.1828 | 0.004353 | 0 | 0 | 0.1952 | 0.1871
0.063 | 0.007769 | 0.1824 | 0.004126 | 0 | 0 | 0.1943 | 0.1865
0.064 | 0.007544 | 0.182 | 0.003915 | 0 | 0 | 0.1934 | 0.1859
0.065 | 0.007329 | 0.1816 | 0.003717 | 0 | 0 | 0.1926 | 0.1853
0.066 | 0.007123 | 0.1812 | 0.003532 | 0 | 0 | 0.1918 | 0.1847
0.067 | 0.006925 | 0.1807 | 0.003359 | 0 | 0 | 0.191 | 0.1841
0.068 | 0.006735 | 0.1803 | 0.003197 | 0 | 0 | 0.1903 | 0.1835
0.069 | 0.006553 | 0.1799 | 0.003045 | 0 | 0 | 0.1895 | 0.183
0.07 | 0.006377 | 0.1795 | 0.002902 | 0 | 0 | 0.1888 | 0.1824
0.071 | 0.006209 | 0.1791 | 0.002768 | 0 | 0 | 0.1881 | 0.1819
0.072 | 0.006048 | 0.1787 | 0.002641 | 0 | 0 | 0.1874 | 0.1813
0.073 | 0.005892 | 0.1783 | 0.002522 | 0 | 0 | 0.1867 | 0.1808
0.074 | 0.005742 | 0.1779 | 0.00241 | 0 | 0 | 0.186 | 0.1803
0.075 | 0.005598 | 0.1775 | 0.002305 | 0 | 0 | 0.1854 | 0.1798
0.076 | 0.005459 | 0.1771 | 0.002205 | 0 | 0 | 0.1847 | 0.1793
0.077 | 0.005326 | 0.1767 | 0.002111 | 0 | 0 | 0.1841 | 0.1788
0.078 | 0.005197 | 0.1763 | 0.002022 | 0 | 0 | 0.1835 | 0.1783
0.079 | 0.005072 | 0.1759 | 0.001938 | 0 | 0 | 0.1829 | 0.1778
0.08 | 0.004952 | 0.1755 | 0.001858 | 0 | 0 | 0.1823 | 0.1773
0.081 | 0.004836 | 0.1751 | 0.001783 | 0 | 0 | 0.1817 | 0.1769
0.082 | 0.004724 | 0.1747 | 0.001711 | 0 | 0 | 0.1811 | 0.1764
0.083 | 0.004616 | 0.1743 | 0.001643 | 0 | 0 | 0.1805 | 0.1759
0.084 | 0.004511 | 0.1739 | 0.001579 | 0 | 0 | 0.18 | 0.1755
0.085 | 0.00441 | 0.1735 | 0.001518 | 0 | 0 | 0.1794 | 0.175
0.086 | 0.004312 | 0.1731 | 0.00146 | 0 | 0 | 0.1789 | 0.1746
0.087 | 0.004218 | 0.1727 | 0.001404 | 0 | 0 | 0.1783 | 0.1741
0.088 | 0.004126 | 0.1723 | 0.001352 | 0 | 0 | 0.1778 | 0.1737
0.089 | 0.004038 | 0.1719 | 0.001302 | 0 | 0 | 0.1773 | 0.1732
0.09 | 0.003952 | 0.1716 | 0.001254 | 0 | 0 | 0.1768 | 0.1728
0.091 | 0.003869 | 0.1712 | 0.001209 | 0 | 0 | 0.1763 | 0.1724
0.092 | 0.003788 | 0.1708 | 0.001166 | 0 | 0 | 0.1757 | 0.172
0.093 | 0.00371 | 0.1704 | 0.001125 | 0 | 0 | 0.1752 | 0.1715
0.094 | 0.003634 | 0.17 | 0.001085 | 0 | 0 | 0.1748 | 0.1711
0.095 | 0.003561 | 0.1697 | 0.001048 | 0 | 0 | 0.1743 | 0.1707
0.096 | 0.003489 | 0.1693 | 0.001012 | 0 | 0 | 0.1738 | 0.1703
0.097 | 0.00342 | 0.1689 | 0.0009774 | 0 | 0 | 0.1733 | 0.1699
0.098 | 0.003353 | 0.1685 | 0.0009446 | 0 | 0 | 0.1728 | 0.1695
0.099 | 0.003288 | 0.1682 | 0.0009133 | 0 | 0 | 0.1724 | 0.1691
0.1 | 0.003224 | 0.1678 | 0.0008833 | 0 | 0 | 0.1719 | 0.1687
0.101 | 0.003163 | 0.1674 | 0.0008545 | 0 | 0 | 0.1715 | 0.1683
0.102 | 0.003103 | 0.1671 | 0.000827 | 0 | 0 | 0.171 | 0.1679
0.103 | 0.003045 | 0.1667 | 0.0008007 | 0 | 0 | 0.1706 | 0.1675
0.104 | 0.002988 | 0.1664 | 0.0007754 | 0 | 0 | 0.1701 | 0.1671
0.105 | 0.002933 | 0.166 | 0.0007511 | 0 | 0 | 0.1697 | 0.1668
0.106 | 0.00288 | 0.1656 | 0.0007279 | 0 | 0 | 0.1693 | 0.1664
0.107 | 0.002828 | 0.1653 | 0.0007056 | 0 | 0 | 0.1688 | 0.166
0.108 | 0.002777 | 0.1649 | 0.0006841 | 0 | 0 | 0.1684 | 0.1656
0.109 | 0.002728 | 0.1646 | 0.0006635 | 0 | 0 | 0.168 | 0.1653
0.11 | 0.00268 | 0.1642 | 0.0006438 | 0 | 0 | 0.1676 | 0.1649
0.111 | 0.002633 | 0.1639 | 0.0006247 | 0 | 0 | 0.1672 | 0.1645
0.112 | 0.002588 | 0.1636 | 0.0006065 | 0 | 0 | 0.1667 | 0.1642
0.113 | 0.002543 | 0.1632 | 0.0005889 | 0 | 0 | 0.1663 | 0.1638
0.114 | 0.0025 | 0.1629 | 0.000572 | 0 | 0 | 0.1659 | 0.1634
0.115 | 0.002458 | 0.1625 | 0.0005557 | 0 | 0 | 0.1655 | 0.1631
0.116 | 0.002417 | 0.1622 | 0.00054 | 0 | 0 | 0.1651 | 0.1627
0.117 | 0.002376 | 0.1619 | 0.0005249 | 0 | 0 | 0.1648 | 0.1624
0.118 | 0.002337 | 0.1615 | 0.0005103 | 0 | 0 | 0.1644 | 0.162
0.119 | 0.002299 | 0.1612 | 0.0004963 | 0 | 0 | 0.164 | 0.1617
0.12 | 0.002262 | 0.1609 | 0.0004827 | 0 | 0 | 0.1636 | 0.1613
0.121 | 0.002226 | 0.1605 | 0.0004697 | 0 | 0 | 0.1632 | 0.161
0.122 | 0.00219 | 0.1602 | 0.0004571 | 0 | 0 | 0.1629 | 0.1607
0.123 | 0.002155 | 0.1599 | 0.0004449 | 0 | 0 | 0.1625 | 0.1603
0.124 | 0.002121 | 0.1596 | 0.0004332 | 0 | 0 | 0.1621 | 0.16
0.125 | 0.002088 | 0.1592 | 0.0004219 | 0 | 0 | 0.1617 | 0.1597
0.126 | 0.002056 | 0.1589 | 0.0004109 | 0 | 0 | 0.1614 | 0.1593
0.127 | 0.002025 | 0.1586 | 0.0004004 | 0 | 0 | 0.161 | 0.159
0.128 | 0.001994 | 0.1583 | 0.0003902 | 0 | 0 | 0.1607 | 0.1587
0.129 | 0.001964 | 0.158 | 0.0003803 | 0 | 0 | 0.1603 | 0.1583
0.13 | 0.001934 | 0.1577 | 0.0003707 | 0 | 0 | 0.16 | 0.158
0.131 | 0.001905 | 0.1573 | 0.0003615 | 0 | 0 | 0.1596 | 0.1577
0.132 | 0.001877 | 0.157 | 0.0003526 | 0 | 0 | 0.1593 | 0.1574
0.133 | 0.00185 | 0.1567 | 0.0003439 | 0 | 0 | 0.1589 | 0.1571
0.134 | 0.001823 | 0.1564 | 0.0003355 | 0 | 0 | 0.1586 | 0.1568
0.135 | 0.001796 | 0.1561 | 0.0003274 | 0 | 0 | 0.1582 | 0.1564
0.136 | 0.00177 | 0.1558 | 0.0003196 | 0 | 0 | 0.1579 | 0.1561
0.137 | 0.001745 | 0.1555 | 0.000312 | 0 | 0 | 0.1576 | 0.1558
0.138 | 0.00172 | 0.1552 | 0.0003046 | 0 | 0 | 0.1572 | 0.1555
0.139 | 0.001696 | 0.1549 | 0.0002975 | 0 | 0 | 0.1569 | 0.1552
0.14 | 0.001672 | 0.1546 | 0.0002906 | 0 | 0 | 0.1566 | 0.1549
0.145 | 0.001561 | 0.1531 | 0.000259 | 0 | 0 | 0.155 | 0.1534
0.15 | 0.00146 | 0.1517 | 0.0002317 | 0 | 0 | 0.1534 | 0.152
0.155 | 0.001369 | 0.1503 | 0.0002082 | 0 | 0 | 0.1519 | 0.1505
0.16 | 0.001286 | 0.149 | 0.0001877 | 0 | 0 | 0.1505 | 0.1492
0.165 | 0.00121 | 0.1477 | 0.0001698 | 0 | 0 | 0.149 | 0.1478
0.17 | 0.001141 | 0.1464 | 0.000154 | 0 | 0 | 0.1477 | 0.1465
0.175 | 0.001077 | 0.1451 | 0.0001402 | 0 | 0 | 0.1463 | 0.1453
0.18 | 0.001019 | 0.1439 | 0.0001279 | 0 | 0 | 0.145 | 0.144
0.185 | 0.000965 | 0.1427 | 0.0001171 | 0 | 0 | 0.1438 | 0.1428
0.19 | 0.0009154 | 0.1415 | 0.0001074 | 0 | 0 | 0.1425 | 0.1416
0.195 | 0.0008695 | 0.1404 | 9.88e-05 | 0 | 0 | 0.1413 | 0.1405
0.2 | 0.0008269 | 0.1393 | 9.107e-05 | 0 | 0 | 0.1402 | 0.1394
0.21 | 0.0007507 | 0.1371 | 7.789e-05 | 0 | 0 | 0.1379 | 0.1372
0.22 | 0.0006845 | 0.1351 | 6.714e-05 | 0 | 0 | 0.1358 | 0.1351
0.23 | 0.0006267 | 0.1331 | 5.831e-05 | 0 | 0 | 0.1338 | 0.1331
0.24 | 0.0005759 | 0.1312 | 5.097e-05 | 0 | 0 | 0.1318 | 0.1312
0.25 | 0.0005311 | 0.1294 | 4.484e-05 | 0 | 0 | 0.13 | 0.1294
0.26 | 0.0004912 | 0.1276 | 3.966e-05 | 0 | 0 | 0.1282 | 0.1277
0.28 | 0.0004239 | 0.1243 | 3.153e-05 | 0 | 0 | 0.1248 | 0.1244
0.3 | 0.0003695 | 0.1213 | 2.552e-05 | 0 | 0 | 0.1217 | 0.1213
0.32 | 0.0003249 | 0.1184 | 2.1e-05 | 0 | 0 | 0.1188 | 0.1185
0.34 | 0.0002879 | 0.1158 | 1.751e-05 | 0 | 0 | 0.1161 | 0.1158
0.36 | 0.0002569 | 0.1133 | 1.479e-05 | 0 | 0 | 0.1136 | 0.1133
0.38 | 0.0002307 | 0.1109 | 1.263e-05 | 0 | 0 | 0.1112 | 0.111
0.4 | 0.0002082 | 0.1087 | 1.089e-05 | 0 | 0 | 0.109 | 0.1087
0.42 | 0.0001889 | 0.1066 | 9.472e-06 | 0 | 0 | 0.1068 | 0.1067
0.44 | 0.0001722 | 0.1047 | 8.307e-06 | 0 | 0 | 0.1049 | 0.1047
0.46 | 0.0001575 | 0.1028 | 7.341e-06 | 0 | 0 | 0.103 | 0.1028
0.48 | 0.0001447 | 0.101 | 6.535e-06 | 0 | 0 | 0.1012 | 0.101
0.5 | 0.0001334 | 0.09933 | 5.857e-06 | 0 | 0 | 0.09947 | 0.09934
0.55 | 0.0001103 | 0.0954 | 4.566e-06 | 0 | 0 | 0.09552 | 0.09541
0.6 | 9.268e-05 | 0.09188 | 3.642e-06 | 0 | 0 | 0.09198 | 0.09189
0.65 | 7.898e-05 | 0.08871 | 2.95e-06 | 0 | 0 | 0.08879 | 0.08871
0.7 | 6.81e-05 | 0.08582 | 2.444e-06 | 0 | 0 | 0.0859 | 0.08583
0.75 | 5.932e-05 | 0.08317 | 2.085e-06 | 0 | 0 | 0.08324 | 0.08318
0.8 | 5.214e-05 | 0.08072 | 1.837e-06 | 0 | 0 | 0.08078 | 0.08072
0.85 | 4.619e-05 | 0.07844 | 1.664e-06 | 0 | 0 | 0.07849 | 0.07845
0.9 | 4.12e-05 | 0.07634 | 1.509e-06 | 0 | 0 | 0.07638 | 0.07634
0.95 | 3.698e-05 | 0.07439 | 1.338e-06 | 0 | 0 | 0.07442 | 0.07439
1 | 3.338e-05 | 0.07259 | 1.142e-06 | 0 | 0 | 0.07263 | 0.07259
2 | 8.347e-06 | 0.05035 | 3.277e-07 | 0.000288 | 0 | 0.05064 | 0.05064
3 | 3.71e-06 | 0.03961 | 1.839e-07 | 0.0008247 | 1.386e-05 | 0.04045 | 0.04045
4 | 2.087e-06 | 0.03304 | 1.265e-07 | 0.001339 | 5.658e-05 | 0.03444 | 0.03444
5 | 1.336e-06 | 0.02854 | 9.614e-08 | 0.001797 | 0.0001128 | 0.03045 | 0.03045
6 | 9.277e-07 | 0.02522 | 7.738e-08 | 0.00221 | 0.0001732 | 0.02761 | 0.0276
7 | 6.813e-07 | 0.02267 | 6.471e-08 | 0.002577 | 0.0002335 | 0.02548 | 0.02548
8 | 5.218e-07 | 0.02063 | 5.556e-08 | 0.002909 | 0.0002921 | 0.02383 | 0.02383
9 | 4.122e-07 | 0.01897 | 4.869e-08 | 0.00321 | 0.0003482 | 0.02252 | 0.02252
10 | 3.339e-07 | 0.01757 | 4.333e-08 | 0.003483 | 0.0004018 | 0.02145 | 0.02145
11 | 2.76e-07 | 0.01639 | 3.899e-08 | 0.003732 | 0.0004526 | 0.02057 | 0.02057
12 | 2.319e-07 | 0.01537 | 3.546e-08 | 0.003961 | 0.0005006 | 0.01983 | 0.01983
13 | 1.976e-07 | 0.01448 | 3.251e-08 | 0.004174 | 0.000546 | 0.0192 | 0.0192
14 | 1.704e-07 | 0.0137 | 3.001e-08 | 0.004372 | 0.0005896 | 0.01866 | 0.01866
15 | 1.484e-07 | 0.01301 | 2.787e-08 | 0.004555 | 0.0006308 | 0.01819 | 0.01819
16 | 1.304e-07 | 0.01238 | 2.601e-08 | 0.004732 | 0.0006698 | 0.01778 | 0.01778
17 | 1.155e-07 | 0.01182 | 2.439e-08 | 0.004897 | 0.0007073 | 0.01743 | 0.01743
18 | 1.031e-07 | 0.01132 | 2.295e-08 | 0.00505 | 0.0007431 | 0.01711 | 0.01711
19 | 9.249e-08 | 0.01086 | 2.168e-08 | 0.005196 | 0.0007773 | 0.01683 | 0.01683
20 | 8.347e-08 | 0.01044 | 2.053e-08 | 0.005336 | 0.0008101 | 0.01658 | 0.01658

## Water
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.08856 | 0.1774 | 0.5439 | 0 | 0 | 0.8098 | 0.7213
0.021 | 0.08237 | 0.1783 | 0.4646 | 0 | 0 | 0.7253 | 0.643
0.022 | 0.07679 | 0.1792 | 0.3997 | 0 | 0 | 0.6557 | 0.5789
0.023 | 0.07174 | 0.18 | 0.3461 | 0 | 0 | 0.5978 | 0.5261
0.024 | 0.06717 | 0.1806 | 0.3015 | 0 | 0 | 0.5493 | 0.4821
0.025 | 0.06301 | 0.1812 | 0.264 | 0 | 0 | 0.5082 | 0.4452
0.026 | 0.05921 | 0.1817 | 0.2324 | 0 | 0 | 0.4733 | 0.4141
0.027 | 0.05574 | 0.1821 | 0.2055 | 0 | 0 | 0.4433 | 0.3876
0.028 | 0.05256 | 0.1824 | 0.1826 | 0 | 0 | 0.4175 | 0.365
0.029 | 0.04963 | 0.1827 | 0.1628 | 0 | 0 | 0.3951 | 0.3455
0.03 | 0.04694 | 0.1829 | 0.1458 | 0 | 0 | 0.3756 | 0.3286
0.031 | 0.04445 | 0.183 | 0.1309 | 0 | 0 | 0.3584 | 0.314
0.032 | 0.04215 | 0.1831 | 0.118 | 0 | 0 | 0.3433 | 0.3012
0.033 | 0.04002 | 0.1832 | 0.1067 | 0 | 0 | 0.33 | 0.2899
0.034 | 0.03804 | 0.1832 | 0.09679 | 0 | 0 | 0.3181 | 0.28
0.035 | 0.03621 | 0.1832 | 0.08802 | 0 | 0 | 0.3075 | 0.2713
0.036 | 0.0345 | 0.1832 | 0.08026 | 0 | 0 | 0.298 | 0.2635
0.037 | 0.03291 | 0.1831 | 0.07336 | 0 | 0 | 0.2894 | 0.2565
0.038 | 0.03143 | 0.183 | 0.06722 | 0 | 0 | 0.2817 | 0.2502
0.039 | 0.03004 | 0.1829 | 0.06172 | 0 | 0 | 0.2747 | 0.2446
0.04 | 0.02874 | 0.1827 | 0.0568 | 0 | 0 | 0.2683 | 0.2395
0.041 | 0.02753 | 0.1826 | 0.05237 | 0 | 0 | 0.2625 | 0.2349
0.042 | 0.02639 | 0.1824 | 0.04838 | 0 | 0 | 0.2571 | 0.2308
0.043 | 0.02532 | 0.1822 | 0.04478 | 0 | 0 | 0.2523 | 0.2269
0.044 | 0.02431 | 0.1819 | 0.04152 | 0 | 0 | 0.2478 | 0.2235
0.045 | 0.02336 | 0.1817 | 0.03856 | 0 | 0 | 0.2436 | 0.2203
0.046 | 0.02247 | 0.1814 | 0.03587 | 0 | 0 | 0.2398 | 0.2173
0.047 | 0.02162 | 0.1812 | 0.03341 | 0 | 0 | 0.2362 | 0.2146
0.048 | 0.02083 | 0.1809 | 0.03117 | 0 | 0 | 0.2329 | 0.2121
0.049 | 0.02007 | 0.1806 | 0.02913 | 0 | 0 | 0.2298 | 0.2097
0.05 | 0.01936 | 0.1803 | 0.02725 | 0 | 0 | 0.2269 | 0.2076
0.051 | 0.01868 | 0.18 | 0.02553 | 0 | 0 | 0.2242 | 0.2055
0.052 | 0.01804 | 0.1797 | 0.02394 | 0 | 0 | 0.2217 | 0.2037
0.053 | 0.01743 | 0.1794 | 0.02248 | 0 | 0 | 0.2193 | 0.2019
0.054 | 0.01686 | 0.1791 | 0.02114 | 0 | 0 | 0.2171 | 0.2002
0.055 | 0.01631 | 0.1787 | 0.0199 | 0 | 0 | 0.2149 | 0.1986
0.056 | 0.01578 | 0.1784 | 0.01875 | 0 | 0 | 0.2129 | 0.1972
0.057 | 0.01528 | 0.1781 | 0.01768 | 0 | 0 | 0.211 | 0.1957
0.058 | 0.01481 | 0.1777 | 0.01669 | 0 | 0 | 0.2092 | 0.1944
0.059 | 0.01435 | 0.1774 | 0.01578 | 0 | 0 | 0.2075 | 0.1932
0.06 | 0.01392 | 0.177 | 0.01493 | 0 | 0 | 0.2059 | 0.192
0.061 | 0.0135 | 0.1767 | 0.01413 | 0 | 0 | 0.2043 | 0.1908
0.062 | 0.01311 | 0.1763 | 0.01339 | 0 | 0 | 0.2028 | 0.1897
0.063 | 0.01273 | 0.176 | 0.0127 | 0 | 0 | 0.2014 | 0.1887
0.064 | 0.01236 | 0.1756 | 0.01206 | 0 | 0 | 0.2 | 0.1877
0.065 | 0.01202 | 0.1752 | 0.01146 | 0 | 0 | 0.1987 | 0.1867
0.066 | 0.01168 | 0.1749 | 0.01089 | 0 | 0 | 0.1974 | 0.1858
0.067 | 0.01136 | 0.1745 | 0.01037 | 0 | 0 | 0.1962 | 0.1849
0.068 | 0.01106 | 0.1741 | 0.009871 | 0 | 0 | 0.1951 | 0.184
0.069 | 0.01076 | 0.1738 | 0.009406 | 0 | 0 | 0.1939 | 0.1832
0.07 | 0.01048 | 0.1734 | 0.008969 | 0 | 0 | 0.1929 | 0.1824
0.071 | 0.0102 | 0.173 | 0.008559 | 0 | 0 | 0.1918 | 0.1816
0.072 | 0.009942 | 0.1727 | 0.008172 | 0 | 0 | 0.1908 | 0.1808
0.073 | 0.00969 | 0.1723 | 0.007808 | 0 | 0 | 0.1898 | 0.1801
0.074 | 0.009448 | 0.1719 | 0.007465 | 0 | 0 | 0.1888 | 0.1794
0.075 | 0.009214 | 0.1716 | 0.007141 | 0 | 0 | 0.1879 | 0.1787
0.076 | 0.008988 | 0.1712 | 0.006835 | 0 | 0 | 0.187 | 0.178
0.077 | 0.008771 | 0.1708 | 0.006546 | 0 | 0 | 0.1861 | 0.1774
0.078 | 0.008562 | 0.1705 | 0.006273 | 0 | 0 | 0.1853 | 0.1767
0.079 | 0.00836 | 0.1701 | 0.006015 | 0 | 0 | 0.1845 | 0.1761
0.08 | 0.008165 | 0.1697 | 0.00577 | 0 | 0 | 0.1837 | 0.1755
0.081 | 0.007976 | 0.1694 | 0.005538 | 0 | 0 | 0.1829 | 0.1749
0.082 | 0.007794 | 0.169 | 0.005318 | 0 | 0 | 0.1821 | 0.1743
0.083 | 0.007618 | 0.1686 | 0.005109 | 0 | 0 | 0.1814 | 0.1737
0.084 | 0.007448 | 0.1683 | 0.004911 | 0 | 0 | 0.1806 | 0.1732
0.085 | 0.007284 | 0.1679 | 0.004723 | 0 | 0 | 0.1799 | 0.1726
0.086 | 0.007125 | 0.1675 | 0.004544 | 0 | 0 | 0.1792 | 0.1721
0.087 | 0.006971 | 0.1672 | 0.004374 | 0 | 0 | 0.1785 | 0.1716
0.088 | 0.006821 | 0.1668 | 0.004212 | 0 | 0 | 0.1779 | 0.171
0.089 | 0.006677 | 0.1665 | 0.004058 | 0 | 0 | 0.1772 | 0.1705
0.09 | 0.006537 | 0.1661 | 0.003911 | 0 | 0 | 0.1766 | 0.17
0.091 | 0.006402 | 0.1658 | 0.003771 | 0 | 0 | 0.1759 | 0.1695
0.092 | 0.00627 | 0.1654 | 0.003637 | 0 | 0 | 0.1753 | 0.169
0.093 | 0.006143 | 0.165 | 0.00351 | 0 | 0 | 0.1747 | 0.1686
0.094 | 0.006019 | 0.1647 | 0.003388 | 0 | 0 | 0.1741 | 0.1681
0.095 | 0.005899 | 0.1643 | 0.003272 | 0 | 0 | 0.1735 | 0.1676
0.096 | 0.005783 | 0.164 | 0.003161 | 0 | 0 | 0.1729 | 0.1672
0.097 | 0.005669 | 0.1636 | 0.003054 | 0 | 0 | 0.1724 | 0.1667
0.098 | 0.00556 | 0.1633 | 0.002953 | 0 | 0 | 0.1718 | 0.1663
0.099 | 0.005453 | 0.163 | 0.002856 | 0 | 0 | 0.1713 | 0.1658
0.1 | 0.005349 | 0.1626 | 0.002763 | 0 | 0 | 0.1707 | 0.1654
0.101 | 0.005249 | 0.1623 | 0.002673 | 0 | 0 | 0.1702 | 0.1649
0.102 | 0.005151 | 0.1619 | 0.002588 | 0 | 0 | 0.1697 | 0.1645
0.103 | 0.005055 | 0.1616 | 0.002506 | 0 | 0 | 0.1692 | 0.1641
0.104 | 0.004963 | 0.1613 | 0.002428 | 0 | 0 | 0.1686 | 0.1637
0.105 | 0.004872 | 0.1609 | 0.002352 | 0 | 0 | 0.1681 | 0.1633
0.106 | 0.004785 | 0.1606 | 0.00228 | 0 | 0 | 0.1677 | 0.1629
0.107 | 0.004699 | 0.1603 | 0.002211 | 0 | 0 | 0.1672 | 0.1625
0.108 | 0.004616 | 0.1599 | 0.002144 | 0 | 0 | 0.1667 | 0.1621
0.109 | 0.004535 | 0.1596 | 0.00208 | 0 | 0 | 0.1662 | 0.1617
0.11 | 0.004456 | 0.1593 | 0.002018 | 0 | 0 | 0.1657 | 0.1613
0.111 | 0.004379 | 0.1589 | 0.001959 | 0 | 0 | 0.1653 | 0.1609
0.112 | 0.004304 | 0.1586 | 0.001902 | 0 | 0 | 0.1648 | 0.1605
0.113 | 0.004231 | 0.1583 | 0.001847 | 0 | 0 | 0.1644 | 0.1601
0.114 | 0.00416 | 0.158 | 0.001795 | 0 | 0 | 0.1639 | 0.1598
0.115 | 0.004091 | 0.1576 | 0.001744 | 0 | 0 | 0.1635 | 0.1594
0.116 | 0.004023 | 0.1573 | 0.001695 | 0 | 0 | 0.163 | 0.159
0.117 | 0.003957 | 0.157 | 0.001648 | 0 | 0 | 0.1626 | 0.1587
0.118 | 0.003892 | 0.1567 | 0.001602 | 0 | 0 | 0.1622 | 0.1583
0.119 | 0.00383 | 0.1564 | 0.001558 | 0 | 0 | 0.1618 | 0.1579
0.12 | 0.003768 | 0.1561 | 0.001516 | 0 | 0 | 0.1614 | 0.1576
0.121 | 0.003708 | 0.1558 | 0.001475 | 0 | 0 | 0.1609 | 0.1572
0.122 | 0.00365 | 0.1554 | 0.001436 | 0 | 0 | 0.1605 | 0.1569
0.123 | 0.003592 | 0.1551 | 0.001398 | 0 | 0 | 0.1601 | 0.1565
0.124 | 0.003537 | 0.1548 | 0.001362 | 0 | 0 | 0.1597 | 0.1562
0.125 | 0.003482 | 0.1545 | 0.001326 | 0 | 0 | 0.1593 | 0.1559
0.126 | 0.003429 | 0.1542 | 0.001292 | 0 | 0 | 0.1589 | 0.1555
0.127 | 0.003376 | 0.1539 | 0.001259 | 0 | 0 | 0.1586 | 0.1552
0.128 | 0.003325 | 0.1536 | 0.001227 | 0 | 0 | 0.1582 | 0.1548
0.129 | 0.003276 | 0.1533 | 0.001196 | 0 | 0 | 0.1578 | 0.1545
0.13 | 0.003227 | 0.153 | 0.001166 | 0 | 0 | 0.1574 | 0.1542
0.131 | 0.003179 | 0.1527 | 0.001137 | 0 | 0 | 0.157 | 0.1539
0.132 | 0.003133 | 0.1524 | 0.001109 | 0 | 0 | 0.1567 | 0.1535
0.133 | 0.003087 | 0.1521 | 0.001082 | 0 | 0 | 0.1563 | 0.1532
0.134 | 0.003042 | 0.1518 | 0.001056 | 0 | 0 | 0.1559 | 0.1529
0.135 | 0.002999 | 0.1516 | 0.001031 | 0 | 0 | 0.1556 | 0.1526
0.136 | 0.002956 | 0.1513 | 0.001006 | 0 | 0 | 0.1552 | 0.1523
0.137 | 0.002914 | 0.151 | 0.0009823 | 0 | 0 | 0.1549 | 0.152
0.138 | 0.002873 | 0.1507 | 0.0009592 | 0 | 0 | 0.1545 | 0.1516
0.139 | 0.002833 | 0.1504 | 0.0009369 | 0 | 0 | 0.1542 | 0.1513
0.14 | 0.002794 | 0.1501 | 0.0009152 | 0 | 0 | 0.1538 | 0.151
0.145 | 0.002609 | 0.1487 | 0.0008161 | 0 | 0 | 0.1521 | 0.1495
0.15 | 0.002442 | 0.1474 | 0.0007308 | 0 | 0 | 0.1505 | 0.1481
0.155 | 0.00229 | 0.146 | 0.0006568 | 0 | 0 | 0.149 | 0.1467
0.16 | 0.002152 | 0.1447 | 0.0005924 | 0 | 0 | 0.1475 | 0.1453
0.165 | 0.002026 | 0.1434 | 0.0005362 | 0 | 0 | 0.146 | 0.144
0.17 | 0.001911 | 0.1422 | 0.0004868 | 0 | 0 | 0.1446 | 0.1427
0.175 | 0.001805 | 0.141 | 0.0004433 | 0 | 0 | 0.1432 | 0.1414
0.18 | 0.001708 | 0.1398 | 0.0004048 | 0 | 0 | 0.1419 | 0.1402
0.185 | 0.001618 | 0.1387 | 0.0003706 | 0 | 0 | 0.1406 | 0.139
0.19 | 0.001535 | 0.1375 | 0.0003402 | 0 | 0 | 0.1394 | 0.1379
0.195 | 0.001459 | 0.1364 | 0.0003131 | 0 | 0 | 0.1382 | 0.1367
0.2 | 0.001388 | 0.1353 | 0.0002887 | 0 | 0 | 0.137 | 0.1356
0.21 | 0.00126 | 0.1333 | 0.0002472 | 0 | 0 | 0.1348 | 0.1335
0.22 | 0.00115 | 0.1313 | 0.0002133 | 0 | 0 | 0.1326 | 0.1315
0.23 | 0.001053 | 0.1294 | 0.0001854 | 0 | 0 | 0.1306 | 0.1295
0.24 | 0.0009677 | 0.1275 | 0.0001623 | 0 | 0 | 0.1287 | 0.1277
0.25 | 0.0008925 | 0.1258 | 0.0001429 | 0 | 0 | 0.1268 | 0.1259
0.26 | 0.0008257 | 0.1241 | 0.0001265 | 0 | 0 | 0.125 | 0.1242
0.28 | 0.0007128 | 0.1209 | 0.0001007 | 0 | 0 | 0.1217 | 0.121
0.3 | 0.0006215 | 0.1179 | 8.16e-05 | 0 | 0 | 0.1186 | 0.118
0.32 | 0.0005467 | 0.1152 | 6.718e-05 | 0 | 0 | 0.1158 | 0.1152
0.34 | 0.0004846 | 0.1126 | 5.608e-05 | 0 | 0 | 0.1131 | 0.1126
0.36 | 0.0004325 | 0.1102 | 4.74e-05 | 0 | 0 | 0.1106 | 0.1102
0.38 | 0.0003883 | 0.1079 | 4.049e-05 | 0 | 0 | 0.1083 | 0.1079
0.4 | 0.0003506 | 0.1058 | 3.493e-05 | 0 | 0 | 0.1061 | 0.1058
0.42 | 0.0003181 | 0.1037 | 3.041e-05 | 0 | 0 | 0.1041 | 0.1038
0.44 | 0.0002899 | 0.1018 | 2.668e-05 | 0 | 0 | 0.1021 | 0.1018
0.46 | 0.0002653 | 0.1 | 2.359e-05 | 0 | 0 | 0.1003 | 0.1
0.48 | 0.0002437 | 0.09828 | 2.101e-05 | 0 | 0 | 0.09854 | 0.0983
0.5 | 0.0002247 | 0.09663 | 1.883e-05 | 0 | 0 | 0.09687 | 0.09665
0.55 | 0.0001857 | 0.09281 | 1.468e-05 | 0 | 0 | 0.09302 | 0.09283
0.6 | 0.0001561 | 0.08939 | 1.173e-05 | 0 | 0 | 0.08956 | 0.0894
0.65 | 0.0001331 | 0.0863 | 9.539e-06 | 0 | 0 | 0.08645 | 0.08631
0.7 | 0.0001148 | 0.0835 | 7.927e-06 | 0 | 0 | 0.08362 | 0.08351
0.75 | 0.0001 | 0.08093 | 6.757e-06 | 0 | 0 | 0.08103 | 0.08093
0.8 | 8.79e-05 | 0.07856 | 5.92e-06 | 0 | 0 | 0.07866 | 0.07857
0.85 | 7.785e-05 | 0.07637 | 5.31e-06 | 0 | 0 | 0.07646 | 0.07638
0.9 | 6.944e-05 | 0.07434 | 4.782e-06 | 0 | 0 | 0.07441 | 0.07434
0.95 | 6.233e-05 | 0.07244 | 4.25e-06 | 0 | 0 | 0.07251 | 0.07244
1 | 5.627e-05 | 0.07066 | 3.681e-06 | 0 | 0 | 0.07072 | 0.07066
2 | 1.407e-05 | 0.04901 | 1.063e-06 | 0.0003908 | 0 | 0.04942 | 0.0494
3 | 6.255e-06 | 0.03855 | 5.937e-07 | 0.001117 | 1.349e-05 | 0.03969 | 0.03968
4 | 3.519e-06 | 0.03216 | 4.075e-07 | 0.001812 | 5.507e-05 | 0.03403 | 0.03402
5 | 2.252e-06 | 0.02777 | 3.09e-07 | 0.002431 | 0.0001097 | 0.03031 | 0.03031
6 | 1.564e-06 | 0.02454 | 2.484e-07 | 0.002987 | 0.0001685 | 0.0277 | 0.0277
7 | 1.149e-06 | 0.02206 | 2.075e-07 | 0.003482 | 0.0002272 | 0.02577 | 0.02577
8 | 8.796e-07 | 0.02008 | 1.78e-07 | 0.003927 | 0.0002843 | 0.02429 | 0.02429
9 | 6.951e-07 | 0.01846 | 1.559e-07 | 0.004334 | 0.0003389 | 0.02313 | 0.02313
10 | 5.63e-07 | 0.0171 | 1.386e-07 | 0.004699 | 0.000391 | 0.02219 | 0.02219
11 | 4.652e-07 | 0.01595 | 1.248e-07 | 0.005034 | 0.0004404 | 0.02142 | 0.02142
12 | 3.909e-07 | 0.01496 | 1.134e-07 | 0.005341 | 0.0004872 | 0.02079 | 0.02079
13 | 3.331e-07 | 0.01409 | 1.039e-07 | 0.005627 | 0.0005315 | 0.02025 | 0.02025
14 | 2.872e-07 | 0.01333 | 9.591e-08 | 0.005893 | 0.0005737 | 0.0198 | 0.0198
15 | 2.502e-07 | 0.01266 | 8.906e-08 | 0.006141 | 0.0006135 | 0.01941 | 0.01941
16 | 2.199e-07 | 0.01205 | 8.311e-08 | 0.006376 | 0.0006519 | 0.01908 | 0.01908
17 | 1.948e-07 | 0.01151 | 7.79e-08 | 0.006596 | 0.0006884 | 0.01879 | 0.01879
18 | 1.737e-07 | 0.01102 | 7.331e-08 | 0.006803 | 0.000723 | 0.01854 | 0.01854
19 | 1.559e-07 | 0.01057 | 6.922e-08 | 0.007 | 0.000756 | 0.01832 | 0.01832
20 | 1.407e-07 | 0.01016 | 6.555e-08 | 0.007186 | 0.0007878 | 0.01813 | 0.01813

## Polystyrene
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.06028 | 0.1753 | 0.2008 | 0 | 0 | 0.4364 | 0.3761
0.021 | 0.05593 | 0.1762 | 0.171 | 0 | 0 | 0.4032 | 0.3473
0.022 | 0.05201 | 0.1771 | 0.1467 | 0 | 0 | 0.3758 | 0.3238
0.023 | 0.04847 | 0.1778 | 0.1267 | 0 | 0 | 0.353 | 0.3045
0.024 | 0.04527 | 0.1783 | 0.1101 | 0 | 0 | 0.3338 | 0.2885
0.025 | 0.04237 | 0.1788 | 0.09625 | 0 | 0 | 0.3174 | 0.2751
0.026 | 0.03973 | 0.1792 | 0.08455 | 0 | 0 | 0.3035 | 0.2638
0.027 | 0.03733 | 0.1795 | 0.07463 | 0 | 0 | 0.2915 | 0.2542
0.028 | 0.03513 | 0.1798 | 0.06617 | 0 | 0 | 0.2811 | 0.2459
0.029 | 0.03312 | 0.1799 | 0.05891 | 0 | 0 | 0.272 | 0.2388
0.03 | 0.03128 | 0.18 | 0.05264 | 0 | 0 | 0.264 | 0.2327
0.031 | 0.02958 | 0.1801 | 0.04722 | 0 | 0 | 0.2569 | 0.2273
0.032 | 0.02802 | 0.1801 | 0.0425 | 0 | 0 | 0.2507 | 0.2226
0.033 | 0.02658 | 0.1801 | 0.03837 | 0 | 0 | 0.2451 | 0.2185
0.034 | 0.02525 | 0.18 | 0.03475 | 0 | 0 | 0.24 | 0.2148
0.035 | 0.02401 | 0.18 | 0.03155 | 0 | 0 | 0.2355 | 0.2115
0.036 | 0.02286 | 0.1798 | 0.02873 | 0 | 0 | 0.2314 | 0.2086
0.037 | 0.02179 | 0.1797 | 0.02623 | 0 | 0 | 0.2277 | 0.2059
0.038 | 0.0208 | 0.1795 | 0.024 | 0 | 0 | 0.2243 | 0.2035
0.039 | 0.01987 | 0.1793 | 0.02201 | 0 | 0 | 0.2212 | 0.2013
0.04 | 0.019 | 0.1791 | 0.02023 | 0 | 0 | 0.2183 | 0.1993
0.041 | 0.01819 | 0.1789 | 0.01864 | 0 | 0 | 0.2157 | 0.1975
0.042 | 0.01742 | 0.1786 | 0.0172 | 0 | 0 | 0.2132 | 0.1958
0.043 | 0.01671 | 0.1784 | 0.0159 | 0 | 0 | 0.211 | 0.1943
0.044 | 0.01604 | 0.1781 | 0.01473 | 0 | 0 | 0.2089 | 0.1928
0.045 | 0.0154 | 0.1778 | 0.01367 | 0 | 0 | 0.2069 | 0.1915
0.046 | 0.01481 | 0.1775 | 0.0127 | 0 | 0 | 0.205 | 0.1902
0.047 | 0.01424 | 0.1772 | 0.01182 | 0 | 0 | 0.2033 | 0.189
0.048 | 0.01371 | 0.1769 | 0.01102 | 0 | 0 | 0.2016 | 0.1879
0.049 | 0.01321 | 0.1766 | 0.01029 | 0 | 0 | 0.2001 | 0.1868
0.05 | 0.01273 | 0.1762 | 0.009618 | 0 | 0 | 0.1986 | 0.1858
0.051 | 0.01228 | 0.1759 | 0.009003 | 0 | 0 | 0.1972 | 0.1849
0.052 | 0.01185 | 0.1755 | 0.008438 | 0 | 0 | 0.1958 | 0.184
0.053 | 0.01145 | 0.1752 | 0.007918 | 0 | 0 | 0.1946 | 0.1831
0.054 | 0.01106 | 0.1748 | 0.007439 | 0 | 0 | 0.1933 | 0.1823
0.055 | 0.01069 | 0.1745 | 0.006997 | 0 | 0 | 0.1922 | 0.1815
0.056 | 0.01035 | 0.1741 | 0.006588 | 0 | 0 | 0.1911 | 0.1807
0.057 | 0.01001 | 0.1738 | 0.00621 | 0 | 0 | 0.19 | 0.18
0.058 | 0.009697 | 0.1734 | 0.005859 | 0 | 0 | 0.1889 | 0.1792
0.059 | 0.009395 | 0.173 | 0.005534 | 0 | 0 | 0.1879 | 0.1786
0.06 | 0.009107 | 0.1726 | 0.005232 | 0 | 0 | 0.187 | 0.1779
0.061 | 0.008832 | 0.1723 | 0.004951 | 0 | 0 | 0.1861 | 0.1772
0.062 | 0.008569 | 0.1719 | 0.004689 | 0 | 0 | 0.1852 | 0.1766
0.063 | 0.008317 | 0.1715 | 0.004445 | 0 | 0 | 0.1843 | 0.176
0.064 | 0.008076 | 0.1712 | 0.004217 | 0 | 0 | 0.1834 | 0.1754
0.065 | 0.007846 | 0.1708 | 0.004005 | 0 | 0 | 0.1826 | 0.1748
0.066 | 0.007625 | 0.1704 | 0.003805 | 0 | 0 | 0.1818 | 0.1742
0.067 | 0.007413 | 0.17 | 0.003619 | 0 | 0 | 0.1811 | 0.1736
0.068 | 0.00721 | 0.1697 | 0.003444 | 0 | 0 | 0.1803 | 0.1731
0.069 | 0.007015 | 0.1693 | 0.00328 | 0 | 0 | 0.1796 | 0.1726
0.07 | 0.006828 | 0.1689 | 0.003126 | 0 | 0 | 0.1789 | 0.172
0.071 | 0.006648 | 0.1685 | 0.002982 | 0 | 0 | 0.1782 | 0.1715
0.072 | 0.006475 | 0.1681 | 0.002846 | 0 | 0 | 0.1775 | 0.171
0.073 | 0.006308 | 0.1678 | 0.002718 | 0 | 0 | 0.1768 | 0.1705
0.074 | 0.006148 | 0.1674 | 0.002597 | 0 | 0 | 0.1761 | 0.17
0.075 | 0.005994 | 0.167 | 0.002483 | 0 | 0 | 0.1755 | 0.1695
0.076 | 0.005845 | 0.1667 | 0.002376 | 0 | 0 | 0.1749 | 0.169
0.077 | 0.005702 | 0.1663 | 0.002274 | 0 | 0 | 0.1743 | 0.1686
0.078 | 0.005564 | 0.1659 | 0.002178 | 0 | 0 | 0.1736 | 0.1681
0.079 | 0.005431 | 0.1655 | 0.002088 | 0 | 0 | 0.1731 | 0.1676
0.08 | 0.005302 | 0.1652 | 0.002002 | 0 | 0 | 0.1725 | 0.1672
0.081 | 0.005178 | 0.1648 | 0.00192 | 0 | 0 | 0.1719 | 0.1667
0.082 | 0.005058 | 0.1644 | 0.001843 | 0 | 0 | 0.1713 | 0.1663
0.083 | 0.004942 | 0.1641 | 0.00177 | 0 | 0 | 0.1708 | 0.1658
0.084 | 0.00483 | 0.1637 | 0.001701 | 0 | 0 | 0.1702 | 0.1654
0.085 | 0.004722 | 0.1633 | 0.001635 | 0 | 0 | 0.1697 | 0.165
0.086 | 0.004617 | 0.163 | 0.001572 | 0 | 0 | 0.1692 | 0.1645
0.087 | 0.004516 | 0.1626 | 0.001513 | 0 | 0 | 0.1686 | 0.1641
0.088 | 0.004418 | 0.1622 | 0.001456 | 0 | 0 | 0.1681 | 0.1637
0.089 | 0.004323 | 0.1619 | 0.001403 | 0 | 0 | 0.1676 | 0.1633
0.09 | 0.004231 | 0.1615 | 0.001351 | 0 | 0 | 0.1671 | 0.1629
0.091 | 0.004142 | 0.1612 | 0.001302 | 0 | 0 | 0.1666 | 0.1625
0.092 | 0.004056 | 0.1608 | 0.001256 | 0 | 0 | 0.1661 | 0.1621
0.093 | 0.003972 | 0.1605 | 0.001211 | 0 | 0 | 0.1656 | 0.1617
0.094 | 0.003891 | 0.1601 | 0.001169 | 0 | 0 | 0.1652 | 0.1613
0.095 | 0.003813 | 0.1598 | 0.001129 | 0 | 0 | 0.1647 | 0.1609
0.096 | 0.003736 | 0.1594 | 0.00109 | 0 | 0 | 0.1642 | 0.1605
0.097 | 0.003662 | 0.1591 | 0.001053 | 0 | 0 | 0.1638 | 0.1601
0.098 | 0.00359 | 0.1587 | 0.001018 | 0 | 0 | 0.1633 | 0.1597
0.099 | 0.003521 | 0.1584 | 0.0009839 | 0 | 0 | 0.1629 | 0.1594
0.1 | 0.003453 | 0.158 | 0.0009516 | 0 | 0 | 0.1624 | 0.159
0.101 | 0.003387 | 0.1577 | 0.0009206 | 0 | 0 | 0.162 | 0.1586
0.102 | 0.003323 | 0.1573 | 0.000891 | 0 | 0 | 0.1616 | 0.1582
0.103 | 0.003261 | 0.157 | 0.0008626 | 0 | 0 | 0.1611 | 0.1579
0.104 | 0.0032 | 0.1567 | 0.0008353 | 0 | 0 | 0.1607 | 0.1575
0.105 | 0.003141 | 0.1563 | 0.0008092 | 0 | 0 | 0.1603 | 0.1571
0.106 | 0.003084 | 0.156 | 0.0007842 | 0 | 0 | 0.1599 | 0.1568
0.107 | 0.003028 | 0.1557 | 0.0007601 | 0 | 0 | 0.1595 | 0.1564
0.108 | 0.002974 | 0.1553 | 0.000737 | 0 | 0 | 0.1591 | 0.1561
0.109 | 0.002921 | 0.155 | 0.0007149 | 0 | 0 | 0.1587 | 0.1557
0.11 | 0.00287 | 0.1547 | 0.0006935 | 0 | 0 | 0.1583 | 0.1554
0.111 | 0.00282 | 0.1544 | 0.0006731 | 0 | 0 | 0.1579 | 0.155
0.112 | 0.002771 | 0.154 | 0.0006534 | 0 | 0 | 0.1575 | 0.1547
0.113 | 0.002723 | 0.1537 | 0.0006344 | 0 | 0 | 0.1571 | 0.1544
0.114 | 0.002677 | 0.1534 | 0.0006162 | 0 | 0 | 0.1567 | 0.154
0.115 | 0.002632 | 0.1531 | 0.0005986 | 0 | 0 | 0.1563 | 0.1537
0.116 | 0.002588 | 0.1528 | 0.0005817 | 0 | 0 | 0.1559 | 0.1534
0.117 | 0.002545 | 0.1525 | 0.0005655 | 0 | 0 | 0.1556 | 0.153
0.118 | 0.002503 | 0.1521 | 0.0005498 | 0 | 0 | 0.1552 | 0.1527
0.119 | 0.002462 | 0.1518 | 0.0005347 | 0 | 0 | 0.1548 | 0.1524
0.12 | 0.002422 | 0.1515 | 0.0005201 | 0 | 0 | 0.1545 | 0.152
0.121 | 0.002383 | 0.1512 | 0.000506 | 0 | 0 | 0.1541 | 0.1517
0.122 | 0.002345 | 0.1509 | 0.0004924 | 0 | 0 | 0.1538 | 0.1514
0.123 | 0.002308 | 0.1506 | 0.0004794 | 0 | 0 | 0.1534 | 0.1511
0.124 | 0.002272 | 0.1503 | 0.0004667 | 0 | 0 | 0.153 | 0.1508
0.125 | 0.002236 | 0.15 | 0.0004545 | 0 | 0 | 0.1527 | 0.1505
0.126 | 0.002202 | 0.1497 | 0.0004427 | 0 | 0 | 0.1523 | 0.1501
0.127 | 0.002168 | 0.1494 | 0.0004313 | 0 | 0 | 0.152 | 0.1498
0.128 | 0.002135 | 0.1491 | 0.0004203 | 0 | 0 | 0.1517 | 0.1495
0.129 | 0.002103 | 0.1488 | 0.0004097 | 0 | 0 | 0.1513 | 0.1492
0.13 | 0.002071 | 0.1485 | 0.0003994 | 0 | 0 | 0.151 | 0.1489
0.131 | 0.00204 | 0.1482 | 0.0003895 | 0 | 0 | 0.1507 | 0.1486
0.132 | 0.00201 | 0.1479 | 0.0003798 | 0 | 0 | 0.1503 | 0.1483
0.133 | 0.001981 | 0.1476 | 0.0003705 | 0 | 0 | 0.15 | 0.148
0.134 | 0.001952 | 0.1474 | 0.0003615 | 0 | 0 | 0.1497 | 0.1477
0.135 | 0.001923 | 0.1471 | 0.0003528 | 0 | 0 | 0.1494 | 0.1474
0.136 | 0.001896 | 0.1468 | 0.0003443 | 0 | 0 | 0.149 | 0.1471
0.137 | 0.001869 | 0.1465 | 0.0003361 | 0 | 0 | 0.1487 | 0.1468
0.138 | 0.001842 | 0.1462 | 0.0003282 | 0 | 0 | 0.1484 | 0.1466
0.139 | 0.001816 | 0.1459 | 0.0003205 | 0 | 0 | 0.1481 | 0.1463
0.14 | 0.001791 | 0.1457 | 0.000313 | 0 | 0 | 0.1478 | 0.146
0.145 | 0.001672 | 0.1443 | 0.000279 | 0 | 0 | 0.1462 | 0.1446
0.15 | 0.001564 | 0.143 | 0.0002497 | 0 | 0 | 0.1448 | 0.1432
0.155 | 0.001466 | 0.1417 | 0.0002243 | 0 | 0 | 0.1433 | 0.1419
0.16 | 0.001377 | 0.1404 | 0.0002022 | 0 | 0 | 0.142 | 0.1406
0.165 | 0.001296 | 0.1391 | 0.0001829 | 0 | 0 | 0.1406 | 0.1393
0.17 | 0.001222 | 0.1379 | 0.0001659 | 0 | 0 | 0.1393 | 0.1381
0.175 | 0.001154 | 0.1367 | 0.000151 | 0 | 0 | 0.1381 | 0.1369
0.18 | 0.001091 | 0.1356 | 0.0001378 | 0 | 0 | 0.1368 | 0.1357
0.185 | 0.001033 | 0.1345 | 0.0001261 | 0 | 0 | 0.1356 | 0.1346
0.19 | 0.0009803 | 0.1334 | 0.0001157 | 0 | 0 | 0.1345 | 0.1335
0.195 | 0.0009312 | 0.1323 | 0.0001064 | 0 | 0 | 0.1333 | 0.1324
0.2 | 0.0008856 | 0.1312 | 9.812e-05 | 0 | 0 | 0.1322 | 0.1313
0.21 | 0.000804 | 0.1292 | 8.391e-05 | 0 | 0 | 0.1301 | 0.1293
0.22 | 0.0007331 | 0.1273 | 7.234e-05 | 0 | 0 | 0.1281 | 0.1274
0.23 | 0.0006712 | 0.1254 | 6.282e-05 | 0 | 0 | 0.1262 | 0.1255
0.24 | 0.0006168 | 0.1236 | 5.491e-05 | 0 | 0 | 0.1243 | 0.1237
0.25 | 0.0005688 | 0.1219 | 4.83e-05 | 0 | 0 | 0.1226 | 0.122
0.26 | 0.0005261 | 0.1203 | 4.273e-05 | 0 | 0 | 0.1209 | 0.1203
0.28 | 0.000454 | 0.1172 | 3.397e-05 | 0 | 0 | 0.1177 | 0.1172
0.3 | 0.0003957 | 0.1143 | 2.75e-05 | 0 | 0 | 0.1147 | 0.1144
0.32 | 0.000348 | 0.1116 | 2.262e-05 | 0 | 0 | 0.112 | 0.1117
0.34 | 0.0003084 | 0.1091 | 1.887e-05 | 0 | 0 | 0.1095 | 0.1092
0.36 | 0.0002752 | 0.1068 | 1.594e-05 | 0 | 0 | 0.1071 | 0.1068
0.38 | 0.000247 | 0.1046 | 1.361e-05 | 0 | 0 | 0.1048 | 0.1046
0.4 | 0.000223 | 0.1025 | 1.173e-05 | 0 | 0 | 0.1027 | 0.1025
0.42 | 0.0002023 | 0.1005 | 1.02e-05 | 0 | 0 | 0.1007 | 0.1005
0.44 | 0.0001844 | 0.09868 | 8.949e-06 | 0 | 0 | 0.09887 | 0.09869
0.46 | 0.0001687 | 0.09691 | 7.909e-06 | 0 | 0 | 0.09709 | 0.09692
0.48 | 0.000155 | 0.09524 | 7.04e-06 | 0 | 0 | 0.0954 | 0.09524
0.5 | 0.0001429 | 0.09364 | 6.31e-06 | 0 | 0 | 0.09379 | 0.09364
0.55 | 0.0001181 | 0.08994 | 4.919e-06 | 0 | 0 | 0.09006 | 0.08994
0.6 | 9.926e-05 | 0.08662 | 3.924e-06 | 0 | 0 | 0.08672 | 0.08662
0.65 | 8.458e-05 | 0.08363 | 3.179e-06 | 0 | 0 | 0.08372 | 0.08363
0.7 | 7.293e-05 | 0.08091 | 2.633e-06 | 0 | 0 | 0.08098 | 0.08091
0.75 | 6.353e-05 | 0.07841 | 2.246e-06 | 0 | 0 | 0.07847 | 0.07841
0.8 | 5.584e-05 | 0.0761 | 1.979e-06 | 0 | 0 | 0.07615 | 0.0761
0.85 | 4.947e-05 | 0.07395 | 1.792e-06 | 0 | 0 | 0.074 | 0.07395
0.9 | 4.413e-05 | 0.07196 | 1.626e-06 | 0 | 0 | 0.07201 | 0.07196
0.95 | 3.961e-05 | 0.07012 | 1.442e-06 | 0 | 0 | 0.07016 | 0.07012
1 | 3.575e-05 | 0.06843 | 1.23e-06 | 0 | 0 | 0.06847 | 0.06843
2 | 8.939e-06 | 0.04746 | 3.53e-07 | 0.0003021 | 0 | 0.04778 | 0.04777
3 | 3.974e-06 | 0.03734 | 1.981e-07 | 0.0008652 | 1.307e-05 | 0.03823 | 0.03822
4 | 2.235e-06 | 0.03115 | 1.363e-07 | 0.001405 | 5.334e-05 | 0.03261 | 0.03261
5 | 1.43e-06 | 0.0269 | 1.036e-07 | 0.001886 | 0.0001063 | 0.0289 | 0.02889
6 | 9.935e-07 | 0.02378 | 8.336e-08 | 0.002318 | 0.0001632 | 0.02626 | 0.02626
7 | 7.297e-07 | 0.02137 | 6.971e-08 | 0.002704 | 0.0002201 | 0.0243 | 0.02429
8 | 5.589e-07 | 0.01945 | 5.986e-08 | 0.003051 | 0.0002753 | 0.02278 | 0.02278
9 | 4.415e-07 | 0.01788 | 5.246e-08 | 0.003367 | 0.0003283 | 0.02157 | 0.02157
10 | 3.576e-07 | 0.01656 | 4.668e-08 | 0.003654 | 0.0003788 | 0.0206 | 0.0206
11 | 2.956e-07 | 0.01545 | 4.201e-08 | 0.003915 | 0.0004267 | 0.01979 | 0.01979
12 | 2.484e-07 | 0.01449 | 3.82e-08 | 0.004155 | 0.000472 | 0.01912 | 0.01912
13 | 2.116e-07 | 0.01365 | 3.502e-08 | 0.004378 | 0.0005147 | 0.01854 | 0.01854
14 | 1.825e-07 | 0.01291 | 3.233e-08 | 0.004586 | 0.0005558 | 0.01806 | 0.01806
15 | 1.59e-07 | 0.01226 | 3.002e-08 | 0.004778 | 0.0005947 | 0.01763 | 0.01763
16 | 1.397e-07 | 0.01167 | 2.802e-08 | 0.004964 | 0.0006315 | 0.01727 | 0.01727
17 | 1.237e-07 | 0.01115 | 2.627e-08 | 0.005136 | 0.0006667 | 0.01695 | 0.01695
18 | 1.104e-07 | 0.01067 | 2.473e-08 | 0.005297 | 0.0007005 | 0.01667 | 0.01667
19 | 9.906e-08 | 0.01024 | 2.335e-08 | 0.00545 | 0.0007328 | 0.01642 | 0.01642
20 | 8.939e-08 | 0.009838 | 2.212e-08 | 0.005597 | 0.0007637 | 0.0162 | 0.0162

## Acrylic
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.071 | 0.1741 | 0.3263 | 0 | 0 | 0.5714 | 0.5004
0.021 | 0.06594 | 0.1751 | 0.2784 | 0 | 0 | 0.5195 | 0.4535
0.022 | 0.06139 | 0.176 | 0.2393 | 0 | 0 | 0.4766 | 0.4152
0.023 | 0.05728 | 0.1767 | 0.207 | 0 | 0 | 0.441 | 0.3837
0.024 | 0.05356 | 0.1773 | 0.1801 | 0 | 0 | 0.411 | 0.3574
0.025 | 0.05018 | 0.1779 | 0.1576 | 0 | 0 | 0.3856 | 0.3355
0.026 | 0.0471 | 0.1783 | 0.1386 | 0 | 0 | 0.364 | 0.3169
0.027 | 0.04429 | 0.1787 | 0.1225 | 0 | 0 | 0.3454 | 0.3011
0.028 | 0.04172 | 0.1789 | 0.1087 | 0 | 0 | 0.3294 | 0.2877
0.029 | 0.03936 | 0.1792 | 0.09689 | 0 | 0 | 0.3154 | 0.276
0.03 | 0.0372 | 0.1793 | 0.08668 | 0 | 0 | 0.3032 | 0.266
0.031 | 0.0352 | 0.1794 | 0.07783 | 0 | 0 | 0.2925 | 0.2573
0.032 | 0.03336 | 0.1795 | 0.07011 | 0 | 0 | 0.283 | 0.2496
0.033 | 0.03166 | 0.1795 | 0.06336 | 0 | 0 | 0.2745 | 0.2429
0.034 | 0.03008 | 0.1795 | 0.05742 | 0 | 0 | 0.267 | 0.2369
0.035 | 0.02862 | 0.1794 | 0.05219 | 0 | 0 | 0.2602 | 0.2316
0.036 | 0.02726 | 0.1793 | 0.04757 | 0 | 0 | 0.2542 | 0.2269
0.037 | 0.02599 | 0.1792 | 0.04346 | 0 | 0 | 0.2487 | 0.2227
0.038 | 0.02481 | 0.1791 | 0.0398 | 0 | 0 | 0.2437 | 0.2189
0.039 | 0.02371 | 0.1789 | 0.03653 | 0 | 0 | 0.2392 | 0.2155
0.04 | 0.02268 | 0.1788 | 0.0336 | 0 | 0 | 0.235 | 0.2123
0.041 | 0.02172 | 0.1785 | 0.03096 | 0 | 0 | 0.2312 | 0.2095
0.042 | 0.02081 | 0.1783 | 0.02859 | 0 | 0 | 0.2277 | 0.2069
0.043 | 0.01996 | 0.1781 | 0.02645 | 0 | 0 | 0.2245 | 0.2045
0.044 | 0.01916 | 0.1778 | 0.02452 | 0 | 0 | 0.2215 | 0.2024
0.045 | 0.01841 | 0.1776 | 0.02276 | 0 | 0 | 0.2188 | 0.2003
0.046 | 0.0177 | 0.1773 | 0.02117 | 0 | 0 | 0.2162 | 0.1985
0.047 | 0.01703 | 0.177 | 0.01971 | 0 | 0 | 0.2138 | 0.1967
0.048 | 0.0164 | 0.1767 | 0.01838 | 0 | 0 | 0.2115 | 0.1951
0.049 | 0.0158 | 0.1764 | 0.01717 | 0 | 0 | 0.2094 | 0.1936
0.05 | 0.01523 | 0.1761 | 0.01606 | 0 | 0 | 0.2074 | 0.1922
0.051 | 0.0147 | 0.1758 | 0.01504 | 0 | 0 | 0.2055 | 0.1908
0.052 | 0.01419 | 0.1755 | 0.0141 | 0 | 0 | 0.2038 | 0.1896
0.053 | 0.01371 | 0.1751 | 0.01324 | 0 | 0 | 0.2021 | 0.1884
0.054 | 0.01325 | 0.1748 | 0.01244 | 0 | 0 | 0.2005 | 0.1872
0.055 | 0.01281 | 0.1745 | 0.01171 | 0 | 0 | 0.199 | 0.1862
0.056 | 0.0124 | 0.1741 | 0.01103 | 0 | 0 | 0.1975 | 0.1851
0.057 | 0.012 | 0.1738 | 0.0104 | 0 | 0 | 0.1962 | 0.1842
0.058 | 0.01163 | 0.1734 | 0.009818 | 0 | 0 | 0.1948 | 0.1832
0.059 | 0.01127 | 0.173 | 0.009277 | 0 | 0 | 0.1936 | 0.1823
0.06 | 0.01092 | 0.1727 | 0.008773 | 0 | 0 | 0.1924 | 0.1815
0.061 | 0.01059 | 0.1723 | 0.008305 | 0 | 0 | 0.1912 | 0.1806
0.062 | 0.01028 | 0.172 | 0.007869 | 0 | 0 | 0.1901 | 0.1798
0.063 | 0.009981 | 0.1716 | 0.007462 | 0 | 0 | 0.189 | 0.1791
0.064 | 0.009694 | 0.1712 | 0.007082 | 0 | 0 | 0.188 | 0.1783
0.065 | 0.009419 | 0.1709 | 0.006727 | 0 | 0 | 0.187 | 0.1776
0.066 | 0.009156 | 0.1705 | 0.006395 | 0 | 0 | 0.1861 | 0.1769
0.067 | 0.008903 | 0.1701 | 0.006084 | 0 | 0 | 0.1851 | 0.1762
0.068 | 0.008661 | 0.1698 | 0.005792 | 0 | 0 | 0.1842 | 0.1756
0.069 | 0.008428 | 0.1694 | 0.005518 | 0 | 0 | 0.1833 | 0.1749
0.07 | 0.008204 | 0.169 | 0.005261 | 0 | 0 | 0.1825 | 0.1743
0.071 | 0.007989 | 0.1687 | 0.005019 | 0 | 0 | 0.1817 | 0.1737
0.072 | 0.007782 | 0.1683 | 0.004791 | 0 | 0 | 0.1809 | 0.1731
0.073 | 0.007584 | 0.1679 | 0.004577 | 0 | 0 | 0.1801 | 0.1725
0.074 | 0.007392 | 0.1676 | 0.004375 | 0 | 0 | 0.1793 | 0.1719
0.075 | 0.007208 | 0.1672 | 0.004184 | 0 | 0 | 0.1786 | 0.1714
0.076 | 0.00703 | 0.1668 | 0.004004 | 0 | 0 | 0.1779 | 0.1708
0.077 | 0.006859 | 0.1665 | 0.003835 | 0 | 0 | 0.1771 | 0.1703
0.078 | 0.006694 | 0.1661 | 0.003674 | 0 | 0 | 0.1765 | 0.1698
0.079 | 0.006535 | 0.1657 | 0.003522 | 0 | 0 | 0.1758 | 0.1692
0.08 | 0.006381 | 0.1654 | 0.003378 | 0 | 0 | 0.1751 | 0.1687
0.081 | 0.006233 | 0.165 | 0.003242 | 0 | 0 | 0.1745 | 0.1682
0.082 | 0.006089 | 0.1646 | 0.003112 | 0 | 0 | 0.1738 | 0.1677
0.083 | 0.005951 | 0.1643 | 0.00299 | 0 | 0 | 0.1732 | 0.1673
0.084 | 0.005817 | 0.1639 | 0.002873 | 0 | 0 | 0.1726 | 0.1668
0.085 | 0.005688 | 0.1635 | 0.002763 | 0 | 0 | 0.172 | 0.1663
0.086 | 0.005562 | 0.1632 | 0.002658 | 0 | 0 | 0.1714 | 0.1658
0.087 | 0.005441 | 0.1628 | 0.002558 | 0 | 0 | 0.1708 | 0.1654
0.088 | 0.005324 | 0.1625 | 0.002463 | 0 | 0 | 0.1703 | 0.1649
0.089 | 0.00521 | 0.1621 | 0.002372 | 0 | 0 | 0.1697 | 0.1645
0.09 | 0.0051 | 0.1618 | 0.002286 | 0 | 0 | 0.1692 | 0.1641
0.091 | 0.004994 | 0.1614 | 0.002204 | 0 | 0 | 0.1686 | 0.1636
0.092 | 0.00489 | 0.1611 | 0.002125 | 0 | 0 | 0.1681 | 0.1632
0.093 | 0.00479 | 0.1607 | 0.002051 | 0 | 0 | 0.1676 | 0.1628
0.094 | 0.004693 | 0.1604 | 0.001979 | 0 | 0 | 0.167 | 0.1623
0.095 | 0.004599 | 0.16 | 0.001911 | 0 | 0 | 0.1665 | 0.1619
0.096 | 0.004507 | 0.1597 | 0.001846 | 0 | 0 | 0.166 | 0.1615
0.097 | 0.004419 | 0.1593 | 0.001784 | 0 | 0 | 0.1655 | 0.1611
0.098 | 0.004332 | 0.159 | 0.001724 | 0 | 0 | 0.165 | 0.1607
0.099 | 0.004249 | 0.1586 | 0.001667 | 0 | 0 | 0.1646 | 0.1603
0.1 | 0.004167 | 0.1583 | 0.001613 | 0 | 0 | 0.1641 | 0.1599
0.101 | 0.004088 | 0.158 | 0.001561 | 0 | 0 | 0.1636 | 0.1595
0.102 | 0.004011 | 0.1576 | 0.001511 | 0 | 0 | 0.1632 | 0.1591
0.103 | 0.003937 | 0.1573 | 0.001463 | 0 | 0 | 0.1627 | 0.1588
0.104 | 0.003864 | 0.157 | 0.001417 | 0 | 0 | 0.1622 | 0.1584
0.105 | 0.003793 | 0.1566 | 0.001373 | 0 | 0 | 0.1618 | 0.158
0.106 | 0.003725 | 0.1563 | 0.00133 | 0 | 0 | 0.1614 | 0.1576
0.107 | 0.003658 | 0.156 | 0.00129 | 0 | 0 | 0.1609 | 0.1573
0.108 | 0.003592 | 0.1557 | 0.001251 | 0 | 0 | 0.1605 | 0.1569
0.109 | 0.003529 | 0.1553 | 0.001213 | 0 | 0 | 0.1601 | 0.1565
0.11 | 0.003467 | 0.155 | 0.001177 | 0 | 0 | 0.1596 | 0.1562
0.111 | 0.003407 | 0.1547 | 0.001143 | 0 | 0 | 0.1592 | 0.1558
0.112 | 0.003348 | 0.1544 | 0.001109 | 0 | 0 | 0.1588 | 0.1555
0.113 | 0.003291 | 0.154 | 0.001077 | 0 | 0 | 0.1584 | 0.1551
0.114 | 0.003236 | 0.1537 | 0.001046 | 0 | 0 | 0.158 | 0.1548
0.115 | 0.003181 | 0.1534 | 0.001017 | 0 | 0 | 0.1576 | 0.1544
0.116 | 0.003128 | 0.1531 | 0.0009882 | 0 | 0 | 0.1572 | 0.1541
0.117 | 0.003077 | 0.1528 | 0.0009606 | 0 | 0 | 0.1568 | 0.1537
0.118 | 0.003026 | 0.1525 | 0.0009341 | 0 | 0 | 0.1564 | 0.1534
0.119 | 0.002977 | 0.1522 | 0.0009085 | 0 | 0 | 0.156 | 0.1531
0.12 | 0.002929 | 0.1519 | 0.0008838 | 0 | 0 | 0.1557 | 0.1527
0.121 | 0.002882 | 0.1515 | 0.00086 | 0 | 0 | 0.1553 | 0.1524
0.122 | 0.002836 | 0.1512 | 0.000837 | 0 | 0 | 0.1549 | 0.1521
0.123 | 0.002792 | 0.1509 | 0.0008149 | 0 | 0 | 0.1545 | 0.1518
0.124 | 0.002748 | 0.1506 | 0.0007934 | 0 | 0 | 0.1542 | 0.1514
0.125 | 0.002705 | 0.1503 | 0.0007728 | 0 | 0 | 0.1538 | 0.1511
0.126 | 0.002664 | 0.15 | 0.0007528 | 0 | 0 | 0.1535 | 0.1508
0.127 | 0.002623 | 0.1497 | 0.0007335 | 0 | 0 | 0.1531 | 0.1505
0.128 | 0.002583 | 0.1494 | 0.0007149 | 0 | 0 | 0.1527 | 0.1502
0.129 | 0.002544 | 0.1492 | 0.0006968 | 0 | 0 | 0.1524 | 0.1498
0.13 | 0.002506 | 0.1489 | 0.0006794 | 0 | 0 | 0.152 | 0.1495
0.131 | 0.002469 | 0.1486 | 0.0006625 | 0 | 0 | 0.1517 | 0.1492
0.132 | 0.002433 | 0.1483 | 0.0006462 | 0 | 0 | 0.1514 | 0.1489
0.133 | 0.002397 | 0.148 | 0.0006304 | 0 | 0 | 0.151 | 0.1486
0.134 | 0.002362 | 0.1477 | 0.0006151 | 0 | 0 | 0.1507 | 0.1483
0.135 | 0.002328 | 0.1474 | 0.0006003 | 0 | 0 | 0.1503 | 0.148
0.136 | 0.002295 | 0.1471 | 0.0005859 | 0 | 0 | 0.15 | 0.1477
0.137 | 0.002262 | 0.1469 | 0.000572 | 0 | 0 | 0.1497 | 0.1474
0.138 | 0.00223 | 0.1466 | 0.0005586 | 0 | 0 | 0.1494 | 0.1471
0.139 | 0.002199 | 0.1463 | 0.0005455 | 0 | 0 | 0.149 | 0.1468
0.14 | 0.002169 | 0.146 | 0.0005329 | 0 | 0 | 0.1487 | 0.1465
0.145 | 0.002024 | 0.1446 | 0.0004751 | 0 | 0 | 0.1471 | 0.1451
0.15 | 0.001894 | 0.1433 | 0.0004253 | 0 | 0 | 0.1456 | 0.1437
0.155 | 0.001776 | 0.142 | 0.0003822 | 0 | 0 | 0.1442 | 0.1424
0.16 | 0.001669 | 0.1407 | 0.0003447 | 0 | 0 | 0.1427 | 0.1411
0.165 | 0.001571 | 0.1395 | 0.0003119 | 0 | 0 | 0.1414 | 0.1398
0.17 | 0.001481 | 0.1383 | 0.0002831 | 0 | 0 | 0.14 | 0.1386
0.175 | 0.001399 | 0.1371 | 0.0002577 | 0 | 0 | 0.1388 | 0.1374
0.18 | 0.001323 | 0.1359 | 0.0002353 | 0 | 0 | 0.1375 | 0.1362
0.185 | 0.001253 | 0.1348 | 0.0002154 | 0 | 0 | 0.1363 | 0.135
0.19 | 0.001189 | 0.1337 | 0.0001977 | 0 | 0 | 0.1351 | 0.1339
0.195 | 0.00113 | 0.1326 | 0.0001819 | 0 | 0 | 0.134 | 0.1328
0.2 | 0.001075 | 0.1316 | 0.0001677 | 0 | 0 | 0.1328 | 0.1318
0.21 | 0.0009756 | 0.1296 | 0.0001435 | 0 | 0 | 0.1307 | 0.1297
0.22 | 0.0008898 | 0.1276 | 0.0001238 | 0 | 0 | 0.1286 | 0.1278
0.23 | 0.0008148 | 0.1258 | 0.0001076 | 0 | 0 | 0.1267 | 0.1259
0.24 | 0.0007488 | 0.124 | 9.41e-05 | 0 | 0 | 0.1248 | 0.1241
0.25 | 0.0006906 | 0.1223 | 8.282e-05 | 0 | 0 | 0.1231 | 0.1224
0.26 | 0.0006388 | 0.1206 | 7.33e-05 | 0 | 0 | 0.1213 | 0.1207
0.28 | 0.0005514 | 0.1175 | 5.832e-05 | 0 | 0 | 0.1181 | 0.1176
0.3 | 0.0004807 | 0.1146 | 4.724e-05 | 0 | 0 | 0.1152 | 0.1147
0.32 | 0.0004227 | 0.112 | 3.889e-05 | 0 | 0 | 0.1124 | 0.112
0.34 | 0.0003747 | 0.1094 | 3.245e-05 | 0 | 0 | 0.1099 | 0.1095
0.36 | 0.0003343 | 0.1071 | 2.742e-05 | 0 | 0 | 0.1075 | 0.1071
0.38 | 0.0003002 | 0.1049 | 2.342e-05 | 0 | 0 | 0.1052 | 0.1049
0.4 | 0.000271 | 0.1028 | 2.02e-05 | 0 | 0 | 0.1031 | 0.1028
0.42 | 0.0002459 | 0.1008 | 1.758e-05 | 0 | 0 | 0.1011 | 0.1008
0.44 | 0.0002241 | 0.09897 | 1.542e-05 | 0 | 0 | 0.09921 | 0.09898
0.46 | 0.0002051 | 0.0972 | 1.363e-05 | 0 | 0 | 0.09742 | 0.09721
0.48 | 0.0001884 | 0.09552 | 1.214e-05 | 0 | 0 | 0.09572 | 0.09553
0.5 | 0.0001736 | 0.09392 | 1.088e-05 | 0 | 0 | 0.0941 | 0.09393
0.55 | 0.0001435 | 0.09021 | 8.482e-06 | 0 | 0 | 0.09036 | 0.09022
0.6 | 0.0001206 | 0.08688 | 6.773e-06 | 0 | 0 | 0.08701 | 0.08689
0.65 | 0.0001028 | 0.08388 | 5.499e-06 | 0 | 0 | 0.08399 | 0.08389
0.7 | 8.867e-05 | 0.08115 | 4.565e-06 | 0 | 0 | 0.08124 | 0.08116
0.75 | 7.725e-05 | 0.07865 | 3.892e-06 | 0 | 0 | 0.07873 | 0.07865
0.8 | 6.79e-05 | 0.07634 | 3.417e-06 | 0 | 0 | 0.07641 | 0.07634
0.85 | 6.014e-05 | 0.07419 | 3.076e-06 | 0 | 0 | 0.07426 | 0.07419
0.9 | 5.365e-05 | 0.0722 | 2.778e-06 | 0 | 0 | 0.07226 | 0.07221
0.95 | 4.815e-05 | 0.07036 | 2.467e-06 | 0 | 0 | 0.07041 | 0.07036
1 | 4.346e-05 | 0.06865 | 2.124e-06 | 0 | 0 | 0.0687 | 0.06865
2 | 1.087e-05 | 0.04762 | 6.12e-07 | 0.000336 | 0 | 0.04796 | 0.04795
3 | 4.832e-06 | 0.03746 | 3.425e-07 | 0.0009616 | 1.311e-05 | 0.03844 | 0.03843
4 | 2.718e-06 | 0.03125 | 2.353e-07 | 0.001561 | 5.351e-05 | 0.03286 | 0.03286
5 | 1.739e-06 | 0.02699 | 1.785e-07 | 0.002094 | 0.0001066 | 0.02919 | 0.02919
6 | 1.208e-06 | 0.02385 | 1.436e-07 | 0.002574 | 0.0001638 | 0.02659 | 0.02659
7 | 8.873e-07 | 0.02144 | 1.2e-07 | 0.003001 | 0.0002208 | 0.02466 | 0.02466
8 | 6.795e-07 | 0.01951 | 1.03e-07 | 0.003386 | 0.0002762 | 0.02317 | 0.02317
9 | 5.369e-07 | 0.01793 | 9.021e-08 | 0.003737 | 0.0003293 | 0.022 | 0.022
10 | 4.349e-07 | 0.01662 | 8.023e-08 | 0.004054 | 0.0003799 | 0.02105 | 0.02105
11 | 3.594e-07 | 0.0155 | 7.221e-08 | 0.004343 | 0.000428 | 0.02027 | 0.02027
12 | 3.02e-07 | 0.01454 | 6.565e-08 | 0.004609 | 0.0004734 | 0.01962 | 0.01962
13 | 2.573e-07 | 0.01369 | 6.017e-08 | 0.004856 | 0.0005164 | 0.01906 | 0.01906
14 | 2.219e-07 | 0.01295 | 5.554e-08 | 0.005086 | 0.0005575 | 0.0186 | 0.0186
15 | 1.933e-07 | 0.0123 | 5.157e-08 | 0.0053 | 0.0005964 | 0.0182 | 0.0182
16 | 1.699e-07 | 0.01171 | 4.813e-08 | 0.005504 | 0.0006334 | 0.01785 | 0.01785
17 | 1.505e-07 | 0.01118 | 4.512e-08 | 0.005695 | 0.0006688 | 0.01755 | 0.01755
18 | 1.342e-07 | 0.0107 | 4.246e-08 | 0.005874 | 0.0007026 | 0.01728 | 0.01728
19 | 1.204e-07 | 0.01027 | 4.009e-08 | 0.006043 | 0.0007349 | 0.01705 | 0.01705
20 | 1.087e-07 | 0.00987 | 3.798e-08 | 0.006205 | 0.0007659 | 0.01684 | 0.01684

## Bone 20
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.09494 | 0.1673 | 1.143 | 0 | 0 | 1.405 | 1.31
0.021 | 0.08824 | 0.1684 | 0.9851 | 0 | 0 | 1.242 | 1.153
0.022 | 0.08221 | 0.1693 | 0.8545 | 0 | 0 | 1.106 | 1.024
0.023 | 0.07676 | 0.1701 | 0.7457 | 0 | 0 | 0.9925 | 0.9158
0.024 | 0.07183 | 0.1708 | 0.6543 | 0 | 0 | 0.897 | 0.8251
0.025 | 0.06735 | 0.1714 | 0.5771 | 0 | 0 | 0.8159 | 0.7485
0.026 | 0.06327 | 0.1719 | 0.5114 | 0 | 0 | 0.7466 | 0.6833
0.027 | 0.05955 | 0.1724 | 0.4552 | 0 | 0 | 0.6871 | 0.6275
0.028 | 0.05615 | 0.1727 | 0.4068 | 0 | 0 | 0.6357 | 0.5795
0.029 | 0.05303 | 0.173 | 0.3649 | 0 | 0 | 0.5909 | 0.5379
0.03 | 0.05016 | 0.1732 | 0.3285 | 0 | 0 | 0.5519 | 0.5017
0.031 | 0.04752 | 0.1734 | 0.2967 | 0 | 0 | 0.5177 | 0.4701
0.032 | 0.04509 | 0.1735 | 0.2689 | 0 | 0 | 0.4875 | 0.4424
0.033 | 0.04284 | 0.1736 | 0.2443 | 0 | 0 | 0.4608 | 0.4179
0.034 | 0.04075 | 0.1736 | 0.2226 | 0 | 0 | 0.437 | 0.3963
0.035 | 0.03881 | 0.1737 | 0.2034 | 0 | 0 | 0.4159 | 0.377
0.036 | 0.03701 | 0.1736 | 0.1863 | 0 | 0 | 0.3969 | 0.3599
0.037 | 0.03533 | 0.1736 | 0.171 | 0 | 0 | 0.3799 | 0.3446
0.038 | 0.03376 | 0.1735 | 0.1573 | 0 | 0 | 0.3646 | 0.3308
0.039 | 0.0323 | 0.1734 | 0.145 | 0 | 0 | 0.3507 | 0.3184
0.04 | 0.03093 | 0.1732 | 0.134 | 0 | 0 | 0.3381 | 0.3072
0.041 | 0.02964 | 0.1731 | 0.124 | 0 | 0 | 0.3267 | 0.2971
0.042 | 0.02844 | 0.1729 | 0.115 | 0 | 0 | 0.3163 | 0.2879
0.043 | 0.0273 | 0.1727 | 0.1068 | 0 | 0 | 0.3068 | 0.2795
0.044 | 0.02624 | 0.1725 | 0.09934 | 0 | 0 | 0.2981 | 0.2718
0.045 | 0.02523 | 0.1723 | 0.09256 | 0 | 0 | 0.2901 | 0.2648
0.046 | 0.02428 | 0.1721 | 0.08637 | 0 | 0 | 0.2827 | 0.2584
0.047 | 0.02339 | 0.1718 | 0.08072 | 0 | 0 | 0.2759 | 0.2525
0.048 | 0.02254 | 0.1716 | 0.07554 | 0 | 0 | 0.2696 | 0.2471
0.049 | 0.02174 | 0.1713 | 0.07078 | 0 | 0 | 0.2638 | 0.2421
0.05 | 0.02098 | 0.171 | 0.06641 | 0 | 0 | 0.2584 | 0.2374
0.051 | 0.02025 | 0.1707 | 0.06239 | 0 | 0 | 0.2534 | 0.2331
0.052 | 0.01957 | 0.1704 | 0.05868 | 0 | 0 | 0.2487 | 0.2291
0.053 | 0.01892 | 0.1701 | 0.05525 | 0 | 0 | 0.2443 | 0.2254
0.054 | 0.0183 | 0.1698 | 0.05209 | 0 | 0 | 0.2402 | 0.2219
0.055 | 0.01771 | 0.1695 | 0.04915 | 0 | 0 | 0.2364 | 0.2187
0.056 | 0.01715 | 0.1692 | 0.04643 | 0 | 0 | 0.2328 | 0.2157
0.057 | 0.01661 | 0.1689 | 0.0439 | 0 | 0 | 0.2294 | 0.2128
0.058 | 0.0161 | 0.1686 | 0.04155 | 0 | 0 | 0.2262 | 0.2101
0.059 | 0.01562 | 0.1683 | 0.03936 | 0 | 0 | 0.2232 | 0.2076
0.06 | 0.01515 | 0.1679 | 0.03732 | 0 | 0 | 0.2204 | 0.2053
0.061 | 0.0147 | 0.1676 | 0.03541 | 0 | 0 | 0.2177 | 0.203
0.062 | 0.01428 | 0.1673 | 0.03363 | 0 | 0 | 0.2152 | 0.2009
0.063 | 0.01387 | 0.1669 | 0.03197 | 0 | 0 | 0.2128 | 0.1989
0.064 | 0.01348 | 0.1666 | 0.03041 | 0 | 0 | 0.2105 | 0.197
0.065 | 0.0131 | 0.1663 | 0.02895 | 0 | 0 | 0.2083 | 0.1952
0.066 | 0.01274 | 0.1659 | 0.02758 | 0 | 0 | 0.2062 | 0.1935
0.067 | 0.0124 | 0.1656 | 0.02629 | 0 | 0 | 0.2043 | 0.1919
0.068 | 0.01207 | 0.1652 | 0.02508 | 0 | 0 | 0.2024 | 0.1903
0.069 | 0.01175 | 0.1649 | 0.02395 | 0 | 0 | 0.2006 | 0.1888
0.07 | 0.01144 | 0.1646 | 0.02288 | 0 | 0 | 0.1989 | 0.1874
0.071 | 0.01115 | 0.1642 | 0.02187 | 0 | 0 | 0.1972 | 0.1861
0.072 | 0.01086 | 0.1639 | 0.02092 | 0 | 0 | 0.1956 | 0.1848
0.073 | 0.01059 | 0.1635 | 0.02002 | 0 | 0 | 0.1941 | 0.1835
0.074 | 0.01033 | 0.1632 | 0.01917 | 0 | 0 | 0.1927 | 0.1823
0.075 | 0.01007 | 0.1628 | 0.01837 | 0 | 0 | 0.1913 | 0.1812
0.076 | 0.009829 | 0.1625 | 0.01761 | 0 | 0 | 0.1899 | 0.1801
0.077 | 0.009594 | 0.1621 | 0.01689 | 0 | 0 | 0.1886 | 0.179
0.078 | 0.009367 | 0.1618 | 0.01621 | 0 | 0 | 0.1874 | 0.178
0.079 | 0.009147 | 0.1615 | 0.01557 | 0 | 0 | 0.1862 | 0.177
0.08 | 0.008936 | 0.1611 | 0.01496 | 0 | 0 | 0.185 | 0.1761
0.081 | 0.008731 | 0.1608 | 0.01438 | 0 | 0 | 0.1839 | 0.1751
0.082 | 0.008533 | 0.1604 | 0.01383 | 0 | 0 | 0.1828 | 0.1742
0.083 | 0.008342 | 0.1601 | 0.0133 | 0 | 0 | 0.1817 | 0.1734
0.084 | 0.008157 | 0.1597 | 0.01281 | 0 | 0 | 0.1807 | 0.1725
0.085 | 0.007978 | 0.1594 | 0.01233 | 0 | 0 | 0.1797 | 0.1717
0.086 | 0.007805 | 0.1591 | 0.01188 | 0 | 0 | 0.1787 | 0.1709
0.087 | 0.007638 | 0.1587 | 0.01145 | 0 | 0 | 0.1778 | 0.1702
0.088 | 0.007476 | 0.1584 | 0.01104 | 0 | 0 | 0.1769 | 0.1694
0.089 | 0.007318 | 0.158 | 0.01065 | 0 | 0 | 0.176 | 0.1687
0.09 | 0.007166 | 0.1577 | 0.01028 | 0 | 0 | 0.1752 | 0.168
0.091 | 0.007018 | 0.1574 | 0.009924 | 0 | 0 | 0.1743 | 0.1673
0.092 | 0.006875 | 0.157 | 0.009585 | 0 | 0 | 0.1735 | 0.1666
0.093 | 0.006736 | 0.1567 | 0.00926 | 0 | 0 | 0.1727 | 0.166
0.094 | 0.006602 | 0.1564 | 0.00895 | 0 | 0 | 0.1719 | 0.1653
0.095 | 0.006471 | 0.156 | 0.008654 | 0 | 0 | 0.1712 | 0.1647
0.096 | 0.006344 | 0.1557 | 0.00837 | 0 | 0 | 0.1704 | 0.1641
0.097 | 0.006221 | 0.1554 | 0.008098 | 0 | 0 | 0.1697 | 0.1635
0.098 | 0.006101 | 0.1551 | 0.007838 | 0 | 0 | 0.169 | 0.1629
0.099 | 0.005985 | 0.1547 | 0.007589 | 0 | 0 | 0.1683 | 0.1623
0.1 | 0.005872 | 0.1544 | 0.007349 | 0 | 0 | 0.1676 | 0.1618
0.101 | 0.005762 | 0.1541 | 0.00712 | 0 | 0 | 0.167 | 0.1612
0.102 | 0.005655 | 0.1538 | 0.0069 | 0 | 0 | 0.1663 | 0.1607
0.103 | 0.005551 | 0.1534 | 0.006689 | 0 | 0 | 0.1657 | 0.1601
0.104 | 0.00545 | 0.1531 | 0.006487 | 0 | 0 | 0.1651 | 0.1596
0.105 | 0.005351 | 0.1528 | 0.006292 | 0 | 0 | 0.1644 | 0.1591
0.106 | 0.005256 | 0.1525 | 0.006105 | 0 | 0 | 0.1638 | 0.1586
0.107 | 0.005162 | 0.1522 | 0.005925 | 0 | 0 | 0.1633 | 0.1581
0.108 | 0.005072 | 0.1519 | 0.005752 | 0 | 0 | 0.1627 | 0.1576
0.109 | 0.004983 | 0.1515 | 0.005586 | 0 | 0 | 0.1621 | 0.1571
0.11 | 0.004897 | 0.1512 | 0.005426 | 0 | 0 | 0.1616 | 0.1567
0.111 | 0.004813 | 0.1509 | 0.005272 | 0 | 0 | 0.161 | 0.1562
0.112 | 0.004731 | 0.1506 | 0.005123 | 0 | 0 | 0.1605 | 0.1557
0.113 | 0.004652 | 0.1503 | 0.004981 | 0 | 0 | 0.1599 | 0.1553
0.114 | 0.004574 | 0.15 | 0.004843 | 0 | 0 | 0.1594 | 0.1549
0.115 | 0.004498 | 0.1497 | 0.00471 | 0 | 0 | 0.1589 | 0.1544
0.116 | 0.004424 | 0.1494 | 0.004582 | 0 | 0 | 0.1584 | 0.154
0.117 | 0.004352 | 0.1491 | 0.004459 | 0 | 0 | 0.1579 | 0.1536
0.118 | 0.004281 | 0.1488 | 0.00434 | 0 | 0 | 0.1574 | 0.1531
0.119 | 0.004213 | 0.1485 | 0.004225 | 0 | 0 | 0.1569 | 0.1527
0.12 | 0.004146 | 0.1482 | 0.004114 | 0 | 0 | 0.1565 | 0.1523
0.121 | 0.00408 | 0.1479 | 0.004007 | 0 | 0 | 0.156 | 0.1519
0.122 | 0.004016 | 0.1476 | 0.003903 | 0 | 0 | 0.1555 | 0.1515
0.123 | 0.003954 | 0.1473 | 0.003803 | 0 | 0 | 0.1551 | 0.1511
0.124 | 0.003893 | 0.147 | 0.003707 | 0 | 0 | 0.1546 | 0.1508
0.125 | 0.003833 | 0.1468 | 0.003613 | 0 | 0 | 0.1542 | 0.1504
0.126 | 0.003775 | 0.1465 | 0.003523 | 0 | 0 | 0.1538 | 0.15
0.127 | 0.003718 | 0.1462 | 0.003436 | 0 | 0 | 0.1533 | 0.1496
0.128 | 0.003662 | 0.1459 | 0.003351 | 0 | 0 | 0.1529 | 0.1492
0.129 | 0.003607 | 0.1456 | 0.003269 | 0 | 0 | 0.1525 | 0.1489
0.13 | 0.003554 | 0.1453 | 0.00319 | 0 | 0 | 0.1521 | 0.1485
0.131 | 0.003502 | 0.145 | 0.003114 | 0 | 0 | 0.1517 | 0.1482
0.132 | 0.003451 | 0.1448 | 0.003039 | 0 | 0 | 0.1513 | 0.1478
0.133 | 0.003401 | 0.1445 | 0.002967 | 0 | 0 | 0.1509 | 0.1475
0.134 | 0.003352 | 0.1442 | 0.002898 | 0 | 0 | 0.1505 | 0.1471
0.135 | 0.003304 | 0.1439 | 0.00283 | 0 | 0 | 0.1501 | 0.1468
0.136 | 0.003258 | 0.1437 | 0.002765 | 0 | 0 | 0.1497 | 0.1464
0.137 | 0.003212 | 0.1434 | 0.002701 | 0 | 0 | 0.1493 | 0.1461
0.138 | 0.003167 | 0.1431 | 0.00264 | 0 | 0 | 0.1489 | 0.1458
0.139 | 0.003123 | 0.1429 | 0.00258 | 0 | 0 | 0.1486 | 0.1454
0.14 | 0.00308 | 0.1426 | 0.002522 | 0 | 0 | 0.1482 | 0.1451
0.145 | 0.002878 | 0.1413 | 0.002257 | 0 | 0 | 0.1464 | 0.1435
0.15 | 0.002695 | 0.14 | 0.002027 | 0 | 0 | 0.1447 | 0.142
0.155 | 0.002528 | 0.1387 | 0.001828 | 0 | 0 | 0.1431 | 0.1405
0.16 | 0.002377 | 0.1375 | 0.001654 | 0 | 0 | 0.1415 | 0.1391
0.165 | 0.002239 | 0.1363 | 0.001501 | 0 | 0 | 0.14 | 0.1378
0.17 | 0.002112 | 0.1351 | 0.001366 | 0 | 0 | 0.1386 | 0.1365
0.175 | 0.001996 | 0.1339 | 0.001247 | 0 | 0 | 0.1372 | 0.1352
0.18 | 0.001889 | 0.1328 | 0.001142 | 0 | 0 | 0.1359 | 0.134
0.185 | 0.00179 | 0.1317 | 0.001048 | 0 | 0 | 0.1346 | 0.1328
0.19 | 0.001699 | 0.1307 | 0.0009642 | 0 | 0 | 0.1333 | 0.1316
0.195 | 0.001615 | 0.1296 | 0.0008891 | 0 | 0 | 0.1321 | 0.1305
0.2 | 0.001537 | 0.1286 | 0.0008217 | 0 | 0 | 0.131 | 0.1294
0.21 | 0.001397 | 0.1266 | 0.0007062 | 0 | 0 | 0.1287 | 0.1273
0.22 | 0.001275 | 0.1247 | 0.0006115 | 0 | 0 | 0.1266 | 0.1254
0.23 | 0.001168 | 0.1229 | 0.0005333 | 0 | 0 | 0.1246 | 0.1235
0.24 | 0.001074 | 0.1212 | 0.0004681 | 0 | 0 | 0.1227 | 0.1217
0.25 | 0.0009909 | 0.1195 | 0.0004132 | 0 | 0 | 0.1209 | 0.1199
0.26 | 0.0009171 | 0.1179 | 0.0003668 | 0 | 0 | 0.1192 | 0.1183
0.28 | 0.0007922 | 0.1149 | 0.0002934 | 0 | 0 | 0.116 | 0.1152
0.3 | 0.0006911 | 0.1121 | 0.0002388 | 0 | 0 | 0.113 | 0.1123
0.32 | 0.0006082 | 0.1095 | 0.0001974 | 0 | 0 | 0.1103 | 0.1097
0.34 | 0.0005393 | 0.107 | 0.0001653 | 0 | 0 | 0.1077 | 0.1072
0.36 | 0.0004815 | 0.1047 | 0.0001402 | 0 | 0 | 0.1053 | 0.1049
0.38 | 0.0004325 | 0.1026 | 0.0001201 | 0 | 0 | 0.1031 | 0.1027
0.4 | 0.0003906 | 0.1005 | 0.0001039 | 0 | 0 | 0.101 | 0.1006
0.42 | 0.0003544 | 0.0986 | 9.065e-05 | 0 | 0 | 0.09905 | 0.09869
0.44 | 0.0003231 | 0.09678 | 7.972e-05 | 0 | 0 | 0.09719 | 0.09686
0.46 | 0.0002958 | 0.09506 | 7.062e-05 | 0 | 0 | 0.09542 | 0.09513
0.48 | 0.0002717 | 0.09342 | 6.298e-05 | 0 | 0 | 0.09375 | 0.09348
0.5 | 0.0002505 | 0.09185 | 5.652e-05 | 0 | 0 | 0.09216 | 0.09191
0.55 | 0.0002072 | 0.08823 | 4.415e-05 | 0 | 0 | 0.08848 | 0.08827
0.6 | 0.0001742 | 0.08498 | 3.539e-05 | 0 | 0 | 0.08519 | 0.08501
0.65 | 0.0001485 | 0.08204 | 2.894e-05 | 0 | 0 | 0.08222 | 0.08207
0.7 | 0.0001281 | 0.07938 | 2.416e-05 | 0 | 0 | 0.07953 | 0.0794
0.75 | 0.0001116 | 0.07693 | 2.06e-05 | 0 | 0 | 0.07706 | 0.07695
0.8 | 9.809e-05 | 0.07467 | 1.795e-05 | 0 | 0 | 0.07478 | 0.07469
0.85 | 8.689e-05 | 0.07257 | 1.592e-05 | 0 | 0 | 0.07267 | 0.07259
0.9 | 7.751e-05 | 0.07063 | 1.422e-05 | 0 | 0 | 0.07072 | 0.07064
0.95 | 6.958e-05 | 0.06882 | 1.266e-05 | 0 | 0 | 0.0689 | 0.06884
1 | 6.281e-05 | 0.06715 | 1.116e-05 | 0 | 0 | 0.06723 | 0.06716
2 | 1.572e-05 | 0.04658 | 3.207e-06 | 0.0003972 | 0 | 0.047 | 0.04698
3 | 6.986e-06 | 0.03665 | 1.764e-06 | 0.001132 | 1.282e-05 | 0.0378 | 0.03779
4 | 3.93e-06 | 0.03057 | 1.198e-06 | 0.001835 | 5.235e-05 | 0.03246 | 0.03246
5 | 2.515e-06 | 0.0264 | 9.022e-07 | 0.002459 | 0.0001043 | 0.02897 | 0.02896
6 | 1.747e-06 | 0.02333 | 7.217e-07 | 0.003021 | 0.0001602 | 0.02652 | 0.02651
7 | 1.283e-06 | 0.02097 | 6.006e-07 | 0.003521 | 0.000216 | 0.02471 | 0.02471
8 | 9.825e-07 | 0.01909 | 5.139e-07 | 0.003971 | 0.0002702 | 0.02333 | 0.02333
9 | 7.763e-07 | 0.01754 | 4.488e-07 | 0.00438 | 0.0003221 | 0.02225 | 0.02225
10 | 6.288e-07 | 0.01625 | 3.983e-07 | 0.004751 | 0.0003716 | 0.02138 | 0.02138
11 | 5.197e-07 | 0.01516 | 3.579e-07 | 0.005089 | 0.0004186 | 0.02067 | 0.02067
12 | 4.367e-07 | 0.01422 | 3.249e-07 | 0.005398 | 0.0004631 | 0.02008 | 0.02008
13 | 3.721e-07 | 0.01339 | 2.974e-07 | 0.005687 | 0.0005051 | 0.01959 | 0.01959
14 | 3.208e-07 | 0.01267 | 2.742e-07 | 0.005955 | 0.0005453 | 0.01917 | 0.01917
15 | 2.795e-07 | 0.01203 | 2.544e-07 | 0.006204 | 0.0005832 | 0.01882 | 0.01882
16 | 2.456e-07 | 0.01146 | 2.372e-07 | 0.006442 | 0.0006195 | 0.01852 | 0.01852
17 | 2.176e-07 | 0.01094 | 2.221e-07 | 0.006664 | 0.0006541 | 0.01826 | 0.01826
18 | 1.941e-07 | 0.01047 | 2.089e-07 | 0.006873 | 0.0006871 | 0.01803 | 0.01803
19 | 1.742e-07 | 0.01004 | 1.971e-07 | 0.00707 | 0.0007186 | 0.01783 | 0.01783
20 | 1.572e-07 | 0.009655 | 1.866e-07 | 0.007258 | 0.0007488 | 0.01766 | 0.01766

## Delrin
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.07905 | 0.1702 | 0.4134 | 0 | 0 | 0.6627 | 0.5836
0.021 | 0.07347 | 0.1713 | 0.3529 | 0 | 0 | 0.5977 | 0.5242
0.022 | 0.06844 | 0.1722 | 0.3034 | 0 | 0 | 0.544 | 0.4756
0.023 | 0.06389 | 0.173 | 0.2626 | 0 | 0 | 0.4994 | 0.4356
0.024 | 0.05977 | 0.1736 | 0.2286 | 0 | 0 | 0.462 | 0.4023
0.025 | 0.05603 | 0.1742 | 0.2001 | 0 | 0 | 0.4304 | 0.3743
0.026 | 0.05262 | 0.1747 | 0.1761 | 0 | 0 | 0.4034 | 0.3508
0.027 | 0.0495 | 0.1751 | 0.1557 | 0 | 0 | 0.3803 | 0.3308
0.028 | 0.04665 | 0.1754 | 0.1382 | 0 | 0 | 0.3603 | 0.3136
0.029 | 0.04403 | 0.1757 | 0.1232 | 0 | 0 | 0.343 | 0.2989
0.03 | 0.04162 | 0.1759 | 0.1103 | 0 | 0 | 0.3278 | 0.2862
0.031 | 0.0394 | 0.1761 | 0.09904 | 0 | 0 | 0.3145 | 0.2751
0.032 | 0.03735 | 0.1762 | 0.08924 | 0 | 0 | 0.3028 | 0.2654
0.033 | 0.03545 | 0.1762 | 0.08067 | 0 | 0 | 0.2924 | 0.2569
0.034 | 0.03369 | 0.1763 | 0.07314 | 0 | 0 | 0.2831 | 0.2494
0.035 | 0.03206 | 0.1762 | 0.06649 | 0 | 0 | 0.2748 | 0.2427
0.036 | 0.03054 | 0.1762 | 0.06061 | 0 | 0 | 0.2674 | 0.2368
0.037 | 0.02913 | 0.1761 | 0.05539 | 0 | 0 | 0.2606 | 0.2315
0.038 | 0.02781 | 0.176 | 0.05073 | 0 | 0 | 0.2546 | 0.2268
0.039 | 0.02658 | 0.1759 | 0.04658 | 0 | 0 | 0.2491 | 0.2225
0.04 | 0.02543 | 0.1757 | 0.04285 | 0 | 0 | 0.244 | 0.2186
0.041 | 0.02435 | 0.1756 | 0.0395 | 0 | 0 | 0.2394 | 0.2151
0.042 | 0.02333 | 0.1754 | 0.03649 | 0 | 0 | 0.2352 | 0.2119
0.043 | 0.02238 | 0.1752 | 0.03376 | 0 | 0 | 0.2313 | 0.2089
0.044 | 0.02149 | 0.175 | 0.0313 | 0 | 0 | 0.2277 | 0.2063
0.045 | 0.02065 | 0.1747 | 0.02906 | 0 | 0 | 0.2244 | 0.2038
0.046 | 0.01986 | 0.1745 | 0.02703 | 0 | 0 | 0.2214 | 0.2015
0.047 | 0.01911 | 0.1742 | 0.02517 | 0 | 0 | 0.2185 | 0.1994
0.048 | 0.0184 | 0.1739 | 0.02348 | 0 | 0 | 0.2158 | 0.1974
0.049 | 0.01773 | 0.1737 | 0.02194 | 0 | 0 | 0.2133 | 0.1956
0.05 | 0.0171 | 0.1734 | 0.02052 | 0 | 0 | 0.211 | 0.1939
0.051 | 0.0165 | 0.1731 | 0.01922 | 0 | 0 | 0.2088 | 0.1923
0.052 | 0.01593 | 0.1728 | 0.01802 | 0 | 0 | 0.2067 | 0.1908
0.053 | 0.01539 | 0.1725 | 0.01692 | 0 | 0 | 0.2048 | 0.1894
0.054 | 0.01488 | 0.1721 | 0.01591 | 0 | 0 | 0.2029 | 0.1881
0.055 | 0.01439 | 0.1718 | 0.01497 | 0 | 0 | 0.2012 | 0.1868
0.056 | 0.01393 | 0.1715 | 0.0141 | 0 | 0 | 0.1995 | 0.1856
0.057 | 0.01348 | 0.1712 | 0.0133 | 0 | 0 | 0.1979 | 0.1845
0.058 | 0.01306 | 0.1708 | 0.01256 | 0 | 0 | 0.1964 | 0.1834
0.059 | 0.01266 | 0.1705 | 0.01187 | 0 | 0 | 0.195 | 0.1824
0.06 | 0.01227 | 0.1702 | 0.01122 | 0 | 0 | 0.1937 | 0.1814
0.061 | 0.01191 | 0.1698 | 0.01063 | 0 | 0 | 0.1923 | 0.1804
0.062 | 0.01156 | 0.1695 | 0.01007 | 0 | 0 | 0.1911 | 0.1795
0.063 | 0.01122 | 0.1691 | 0.00955 | 0 | 0 | 0.1899 | 0.1787
0.064 | 0.0109 | 0.1688 | 0.009064 | 0 | 0 | 0.1887 | 0.1778
0.065 | 0.01059 | 0.1684 | 0.008611 | 0 | 0 | 0.1876 | 0.177
0.066 | 0.01029 | 0.1681 | 0.008186 | 0 | 0 | 0.1865 | 0.1762
0.067 | 0.01001 | 0.1677 | 0.007789 | 0 | 0 | 0.1855 | 0.1755
0.068 | 0.009739 | 0.1674 | 0.007416 | 0 | 0 | 0.1845 | 0.1748
0.069 | 0.009478 | 0.167 | 0.007066 | 0 | 0 | 0.1835 | 0.1741
0.07 | 0.009228 | 0.1666 | 0.006737 | 0 | 0 | 0.1826 | 0.1734
0.071 | 0.008987 | 0.1663 | 0.006428 | 0 | 0 | 0.1817 | 0.1727
0.072 | 0.008755 | 0.1659 | 0.006137 | 0 | 0 | 0.1808 | 0.1721
0.073 | 0.008532 | 0.1656 | 0.005863 | 0 | 0 | 0.18 | 0.1714
0.074 | 0.008317 | 0.1652 | 0.005605 | 0 | 0 | 0.1791 | 0.1708
0.075 | 0.00811 | 0.1649 | 0.005361 | 0 | 0 | 0.1783 | 0.1702
0.076 | 0.007911 | 0.1645 | 0.005131 | 0 | 0 | 0.1775 | 0.1696
0.077 | 0.007719 | 0.1641 | 0.004914 | 0 | 0 | 0.1768 | 0.1691
0.078 | 0.007534 | 0.1638 | 0.004708 | 0 | 0 | 0.176 | 0.1685
0.079 | 0.007355 | 0.1634 | 0.004514 | 0 | 0 | 0.1753 | 0.1679
0.08 | 0.007183 | 0.1631 | 0.00433 | 0 | 0 | 0.1746 | 0.1674
0.081 | 0.007016 | 0.1627 | 0.004155 | 0 | 0 | 0.1739 | 0.1669
0.082 | 0.006856 | 0.1624 | 0.00399 | 0 | 0 | 0.1732 | 0.1664
0.083 | 0.0067 | 0.162 | 0.003833 | 0 | 0 | 0.1726 | 0.1659
0.084 | 0.00655 | 0.1617 | 0.003684 | 0 | 0 | 0.1719 | 0.1654
0.085 | 0.006405 | 0.1613 | 0.003543 | 0 | 0 | 0.1713 | 0.1649
0.086 | 0.006264 | 0.161 | 0.003408 | 0 | 0 | 0.1706 | 0.1644
0.087 | 0.006128 | 0.1606 | 0.00328 | 0 | 0 | 0.17 | 0.1639
0.088 | 0.005996 | 0.1603 | 0.003158 | 0 | 0 | 0.1694 | 0.1634
0.089 | 0.005869 | 0.1599 | 0.003043 | 0 | 0 | 0.1688 | 0.163
0.09 | 0.005745 | 0.1596 | 0.002932 | 0 | 0 | 0.1683 | 0.1625
0.091 | 0.005626 | 0.1592 | 0.002827 | 0 | 0 | 0.1677 | 0.1621
0.092 | 0.00551 | 0.1589 | 0.002727 | 0 | 0 | 0.1671 | 0.1616
0.093 | 0.005397 | 0.1586 | 0.002631 | 0 | 0 | 0.1666 | 0.1612
0.094 | 0.005288 | 0.1582 | 0.00254 | 0 | 0 | 0.166 | 0.1608
0.095 | 0.005182 | 0.1579 | 0.002452 | 0 | 0 | 0.1655 | 0.1603
0.096 | 0.005079 | 0.1575 | 0.002369 | 0 | 0 | 0.165 | 0.1599
0.097 | 0.00498 | 0.1572 | 0.002289 | 0 | 0 | 0.1645 | 0.1595
0.098 | 0.004883 | 0.1569 | 0.002213 | 0 | 0 | 0.164 | 0.1591
0.099 | 0.004789 | 0.1565 | 0.00214 | 0 | 0 | 0.1635 | 0.1587
0.1 | 0.004697 | 0.1562 | 0.00207 | 0 | 0 | 0.163 | 0.1583
0.101 | 0.004608 | 0.1559 | 0.002003 | 0 | 0 | 0.1625 | 0.1579
0.102 | 0.004522 | 0.1556 | 0.001939 | 0 | 0 | 0.162 | 0.1575
0.103 | 0.004438 | 0.1552 | 0.001878 | 0 | 0 | 0.1615 | 0.1571
0.104 | 0.004356 | 0.1549 | 0.001819 | 0 | 0 | 0.1611 | 0.1567
0.105 | 0.004277 | 0.1546 | 0.001762 | 0 | 0 | 0.1606 | 0.1563
0.106 | 0.0042 | 0.1543 | 0.001708 | 0 | 0 | 0.1602 | 0.156
0.107 | 0.004124 | 0.1539 | 0.001656 | 0 | 0 | 0.1597 | 0.1556
0.108 | 0.004051 | 0.1536 | 0.001606 | 0 | 0 | 0.1593 | 0.1552
0.109 | 0.00398 | 0.1533 | 0.001558 | 0 | 0 | 0.1588 | 0.1549
0.11 | 0.00391 | 0.153 | 0.001512 | 0 | 0 | 0.1584 | 0.1545
0.111 | 0.003843 | 0.1527 | 0.001467 | 0 | 0 | 0.158 | 0.1541
0.112 | 0.003777 | 0.1523 | 0.001425 | 0 | 0 | 0.1575 | 0.1538
0.113 | 0.003712 | 0.152 | 0.001383 | 0 | 0 | 0.1571 | 0.1534
0.114 | 0.00365 | 0.1517 | 0.001344 | 0 | 0 | 0.1567 | 0.1531
0.115 | 0.003588 | 0.1514 | 0.001306 | 0 | 0 | 0.1563 | 0.1527
0.116 | 0.003529 | 0.1511 | 0.001269 | 0 | 0 | 0.1559 | 0.1524
0.117 | 0.003471 | 0.1508 | 0.001234 | 0 | 0 | 0.1555 | 0.152
0.118 | 0.003414 | 0.1505 | 0.0012 | 0 | 0 | 0.1551 | 0.1517
0.119 | 0.003359 | 0.1502 | 0.001167 | 0 | 0 | 0.1547 | 0.1514
0.12 | 0.003305 | 0.1499 | 0.001135 | 0 | 0 | 0.1543 | 0.151
0.121 | 0.003252 | 0.1496 | 0.001105 | 0 | 0 | 0.1539 | 0.1507
0.122 | 0.0032 | 0.1493 | 0.001075 | 0 | 0 | 0.1536 | 0.1504
0.123 | 0.00315 | 0.149 | 0.001047 | 0 | 0 | 0.1532 | 0.15
0.124 | 0.003101 | 0.1487 | 0.001019 | 0 | 0 | 0.1528 | 0.1497
0.125 | 0.003053 | 0.1484 | 0.0009927 | 0 | 0 | 0.1525 | 0.1494
0.126 | 0.003006 | 0.1481 | 0.0009671 | 0 | 0 | 0.1521 | 0.1491
0.127 | 0.00296 | 0.1478 | 0.0009423 | 0 | 0 | 0.1517 | 0.1488
0.128 | 0.002915 | 0.1475 | 0.0009184 | 0 | 0 | 0.1514 | 0.1484
0.129 | 0.002871 | 0.1472 | 0.0008952 | 0 | 0 | 0.151 | 0.1481
0.13 | 0.002829 | 0.147 | 0.0008729 | 0 | 0 | 0.1507 | 0.1478
0.131 | 0.002787 | 0.1467 | 0.0008512 | 0 | 0 | 0.1503 | 0.1475
0.132 | 0.002746 | 0.1464 | 0.0008303 | 0 | 0 | 0.15 | 0.1472
0.133 | 0.002706 | 0.1461 | 0.00081 | 0 | 0 | 0.1496 | 0.1469
0.134 | 0.002667 | 0.1458 | 0.0007903 | 0 | 0 | 0.1493 | 0.1466
0.135 | 0.002628 | 0.1455 | 0.0007713 | 0 | 0 | 0.1489 | 0.1463
0.136 | 0.002591 | 0.1453 | 0.0007529 | 0 | 0 | 0.1486 | 0.146
0.137 | 0.002554 | 0.145 | 0.0007351 | 0 | 0 | 0.1483 | 0.1457
0.138 | 0.002518 | 0.1447 | 0.0007178 | 0 | 0 | 0.1479 | 0.1454
0.139 | 0.002483 | 0.1444 | 0.000701 | 0 | 0 | 0.1476 | 0.1451
0.14 | 0.002448 | 0.1442 | 0.0006848 | 0 | 0 | 0.1473 | 0.1448
0.145 | 0.002286 | 0.1428 | 0.0006106 | 0 | 0 | 0.1457 | 0.1434
0.15 | 0.002139 | 0.1415 | 0.0005467 | 0 | 0 | 0.1442 | 0.142
0.155 | 0.002006 | 0.1402 | 0.0004913 | 0 | 0 | 0.1427 | 0.1407
0.16 | 0.001885 | 0.139 | 0.0004431 | 0 | 0 | 0.1413 | 0.1394
0.165 | 0.001774 | 0.1377 | 0.000401 | 0 | 0 | 0.1399 | 0.1381
0.17 | 0.001673 | 0.1365 | 0.000364 | 0 | 0 | 0.1386 | 0.1369
0.175 | 0.00158 | 0.1354 | 0.0003314 | 0 | 0 | 0.1373 | 0.1357
0.18 | 0.001495 | 0.1342 | 0.0003026 | 0 | 0 | 0.136 | 0.1346
0.185 | 0.001416 | 0.1331 | 0.0002771 | 0 | 0 | 0.1348 | 0.1334
0.19 | 0.001344 | 0.1321 | 0.0002543 | 0 | 0 | 0.1337 | 0.1323
0.195 | 0.001276 | 0.131 | 0.000234 | 0 | 0 | 0.1325 | 0.1312
0.2 | 0.001214 | 0.13 | 0.0002158 | 0 | 0 | 0.1314 | 0.1302
0.21 | 0.001103 | 0.128 | 0.0001847 | 0 | 0 | 0.1292 | 0.1281
0.22 | 0.001006 | 0.126 | 0.0001593 | 0 | 0 | 0.1272 | 0.1262
0.23 | 0.0009209 | 0.1242 | 0.0001385 | 0 | 0 | 0.1253 | 0.1244
0.24 | 0.0008464 | 0.1225 | 0.0001212 | 0 | 0 | 0.1234 | 0.1226
0.25 | 0.0007806 | 0.1208 | 0.0001067 | 0 | 0 | 0.1217 | 0.1209
0.26 | 0.0007222 | 0.1191 | 9.441e-05 | 0 | 0 | 0.12 | 0.1192
0.28 | 0.0006233 | 0.1161 | 7.513e-05 | 0 | 0 | 0.1168 | 0.1162
0.3 | 0.0005435 | 0.1132 | 6.088e-05 | 0 | 0 | 0.1138 | 0.1133
0.32 | 0.000478 | 0.1106 | 5.012e-05 | 0 | 0 | 0.1111 | 0.1106
0.34 | 0.0004236 | 0.1081 | 4.183e-05 | 0 | 0 | 0.1086 | 0.1081
0.36 | 0.0003781 | 0.1058 | 3.535e-05 | 0 | 0 | 0.1062 | 0.1058
0.38 | 0.0003394 | 0.1036 | 3.019e-05 | 0 | 0 | 0.104 | 0.1036
0.4 | 0.0003065 | 0.1015 | 2.605e-05 | 0 | 0 | 0.1019 | 0.1016
0.42 | 0.0002781 | 0.0996 | 2.267e-05 | 0 | 0 | 0.0999 | 0.09962
0.44 | 0.0002534 | 0.09776 | 1.989e-05 | 0 | 0 | 0.09803 | 0.09778
0.46 | 0.0002319 | 0.09602 | 1.759e-05 | 0 | 0 | 0.09626 | 0.09603
0.48 | 0.000213 | 0.09436 | 1.566e-05 | 0 | 0 | 0.09458 | 0.09437
0.5 | 0.0001964 | 0.09277 | 1.404e-05 | 0 | 0 | 0.09298 | 0.09279
0.55 | 0.0001623 | 0.08911 | 1.094e-05 | 0 | 0 | 0.08928 | 0.08912
0.6 | 0.0001365 | 0.08582 | 8.739e-06 | 0 | 0 | 0.08597 | 0.08583
0.65 | 0.0001163 | 0.08286 | 7.101e-06 | 0 | 0 | 0.08298 | 0.08287
0.7 | 0.0001003 | 0.08016 | 5.897e-06 | 0 | 0 | 0.08027 | 0.08017
0.75 | 8.738e-05 | 0.07769 | 5.028e-06 | 0 | 0 | 0.07779 | 0.0777
0.8 | 7.68e-05 | 0.07542 | 4.41e-06 | 0 | 0 | 0.0755 | 0.07542
0.85 | 6.803e-05 | 0.0733 | 3.963e-06 | 0 | 0 | 0.07338 | 0.07331
0.9 | 6.068e-05 | 0.07134 | 3.574e-06 | 0 | 0 | 0.07141 | 0.07135
0.95 | 5.447e-05 | 0.06952 | 3.175e-06 | 0 | 0 | 0.06958 | 0.06952
1 | 4.917e-05 | 0.06783 | 2.742e-06 | 0 | 0 | 0.06788 | 0.06783
2 | 1.23e-05 | 0.04705 | 7.907e-07 | 0.000362 | 0 | 0.04742 | 0.04741
3 | 5.466e-06 | 0.03701 | 4.421e-07 | 0.001035 | 1.295e-05 | 0.03806 | 0.03806
4 | 3.075e-06 | 0.03087 | 3.036e-07 | 0.00168 | 5.287e-05 | 0.03261 | 0.0326
5 | 1.968e-06 | 0.02666 | 2.303e-07 | 0.002254 | 0.0001053 | 0.02902 | 0.02902
6 | 1.366e-06 | 0.02356 | 1.852e-07 | 0.00277 | 0.0001618 | 0.0265 | 0.0265
7 | 1.004e-06 | 0.02118 | 1.547e-07 | 0.00323 | 0.0002181 | 0.02463 | 0.02463
8 | 7.686e-07 | 0.01927 | 1.328e-07 | 0.003643 | 0.0002729 | 0.02319 | 0.02319
9 | 6.073e-07 | 0.01772 | 1.163e-07 | 0.00402 | 0.0003254 | 0.02207 | 0.02206
10 | 4.919e-07 | 0.01642 | 1.034e-07 | 0.004361 | 0.0003754 | 0.02115 | 0.02115
11 | 4.065e-07 | 0.01531 | 9.307e-08 | 0.004672 | 0.0004228 | 0.02041 | 0.02041
12 | 3.416e-07 | 0.01436 | 8.46e-08 | 0.004957 | 0.0004677 | 0.01979 | 0.01979
13 | 2.911e-07 | 0.01353 | 7.754e-08 | 0.005223 | 0.0005102 | 0.01926 | 0.01926
14 | 2.51e-07 | 0.0128 | 7.156e-08 | 0.00547 | 0.0005508 | 0.01882 | 0.01882
15 | 2.186e-07 | 0.01215 | 6.645e-08 | 0.0057 | 0.0005891 | 0.01844 | 0.01844
16 | 1.921e-07 | 0.01157 | 6.201e-08 | 0.005919 | 0.0006258 | 0.01812 | 0.01812
17 | 1.702e-07 | 0.01105 | 5.813e-08 | 0.006124 | 0.0006608 | 0.01783 | 0.01783
18 | 1.518e-07 | 0.01057 | 5.471e-08 | 0.006316 | 0.0006941 | 0.01758 | 0.01758
19 | 1.363e-07 | 0.01014 | 5.165e-08 | 0.006498 | 0.0007259 | 0.01737 | 0.01737
20 | 1.23e-07 | 0.009751 | 4.892e-08 | 0.006672 | 0.0007565 | 0.01718 | 0.01718

## Bone 50
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.1311 | 0.161 | 2.38 | 0 | 0 | 2.673 | 2.541
0.021 | 0.1219 | 0.1622 | 2.055 | 0 | 0 | 2.339 | 2.218
0.022 | 0.1136 | 0.1633 | 1.786 | 0 | 0 | 2.063 | 1.949
0.023 | 0.1061 | 0.1642 | 1.561 | 0 | 0 | 1.832 | 1.725
0.024 | 0.09939 | 0.165 | 1.372 | 0 | 0 | 1.637 | 1.537
0.025 | 0.09326 | 0.1657 | 1.212 | 0 | 0 | 1.471 | 1.378
0.026 | 0.08768 | 0.1662 | 1.076 | 0 | 0 | 1.33 | 1.242
0.027 | 0.08259 | 0.1668 | 0.9588 | 0 | 0 | 1.208 | 1.126
0.028 | 0.07794 | 0.1672 | 0.8579 | 0 | 0 | 1.103 | 1.025
0.029 | 0.07367 | 0.1676 | 0.7706 | 0 | 0 | 1.012 | 0.9381
0.03 | 0.06975 | 0.1679 | 0.6945 | 0 | 0 | 0.9321 | 0.8624
0.031 | 0.06614 | 0.1681 | 0.628 | 0 | 0 | 0.8623 | 0.7961
0.032 | 0.06281 | 0.1683 | 0.5696 | 0 | 0 | 0.8008 | 0.7379
0.033 | 0.05972 | 0.1685 | 0.5181 | 0 | 0 | 0.7463 | 0.6866
0.034 | 0.05686 | 0.1686 | 0.4726 | 0 | 0 | 0.6981 | 0.6412
0.035 | 0.05421 | 0.1687 | 0.4322 | 0 | 0 | 0.655 | 0.6008
0.036 | 0.05174 | 0.1687 | 0.3961 | 0 | 0 | 0.6166 | 0.5648
0.037 | 0.04944 | 0.1687 | 0.3639 | 0 | 0 | 0.5821 | 0.5327
0.038 | 0.04728 | 0.1687 | 0.3351 | 0 | 0 | 0.5511 | 0.5038
0.039 | 0.04527 | 0.1686 | 0.3092 | 0 | 0 | 0.5231 | 0.4778
0.04 | 0.04339 | 0.1686 | 0.2858 | 0 | 0 | 0.4978 | 0.4544
0.041 | 0.04162 | 0.1685 | 0.2647 | 0 | 0 | 0.4748 | 0.4332
0.042 | 0.03996 | 0.1684 | 0.2456 | 0 | 0 | 0.4539 | 0.414
0.043 | 0.0384 | 0.1682 | 0.2283 | 0 | 0 | 0.4349 | 0.3965
0.044 | 0.03693 | 0.1681 | 0.2125 | 0 | 0 | 0.4175 | 0.3806
0.045 | 0.03554 | 0.1679 | 0.1981 | 0 | 0 | 0.4016 | 0.366
0.046 | 0.03423 | 0.1677 | 0.185 | 0 | 0 | 0.3869 | 0.3527
0.047 | 0.03299 | 0.1675 | 0.173 | 0 | 0 | 0.3735 | 0.3405
0.048 | 0.03182 | 0.1673 | 0.162 | 0 | 0 | 0.3611 | 0.3293
0.049 | 0.0307 | 0.1671 | 0.1519 | 0 | 0 | 0.3497 | 0.319
0.05 | 0.02965 | 0.1669 | 0.1426 | 0 | 0 | 0.3391 | 0.3094
0.051 | 0.02865 | 0.1666 | 0.134 | 0 | 0 | 0.3293 | 0.3006
0.052 | 0.0277 | 0.1664 | 0.1261 | 0 | 0 | 0.3202 | 0.2925
0.053 | 0.02679 | 0.1662 | 0.1188 | 0 | 0 | 0.3117 | 0.2849
0.054 | 0.02593 | 0.1659 | 0.112 | 0 | 0 | 0.3038 | 0.2779
0.055 | 0.02511 | 0.1656 | 0.1058 | 0 | 0 | 0.2965 | 0.2714
0.056 | 0.02433 | 0.1654 | 0.09995 | 0 | 0 | 0.2896 | 0.2653
0.057 | 0.02358 | 0.1651 | 0.09454 | 0 | 0 | 0.2832 | 0.2596
0.058 | 0.02287 | 0.1648 | 0.08952 | 0 | 0 | 0.2772 | 0.2543
0.059 | 0.02219 | 0.1645 | 0.08483 | 0 | 0 | 0.2715 | 0.2493
0.06 | 0.02154 | 0.1642 | 0.08047 | 0 | 0 | 0.2662 | 0.2447
0.061 | 0.02091 | 0.1639 | 0.07639 | 0 | 0 | 0.2612 | 0.2403
0.062 | 0.02032 | 0.1636 | 0.07258 | 0 | 0 | 0.2565 | 0.2362
0.063 | 0.01974 | 0.1633 | 0.06901 | 0 | 0 | 0.2521 | 0.2323
0.064 | 0.01919 | 0.163 | 0.06567 | 0 | 0 | 0.2479 | 0.2287
0.065 | 0.01867 | 0.1627 | 0.06254 | 0 | 0 | 0.2439 | 0.2253
0.066 | 0.01816 | 0.1624 | 0.0596 | 0 | 0 | 0.2402 | 0.222
0.067 | 0.01768 | 0.1621 | 0.05684 | 0 | 0 | 0.2366 | 0.2189
0.068 | 0.01721 | 0.1618 | 0.05424 | 0 | 0 | 0.2332 | 0.216
0.069 | 0.01676 | 0.1615 | 0.0518 | 0 | 0 | 0.23 | 0.2133
0.07 | 0.01633 | 0.1612 | 0.0495 | 0 | 0 | 0.227 | 0.2107
0.071 | 0.01592 | 0.1608 | 0.04733 | 0 | 0 | 0.2241 | 0.2082
0.072 | 0.01552 | 0.1605 | 0.04529 | 0 | 0 | 0.2213 | 0.2058
0.073 | 0.01513 | 0.1602 | 0.04336 | 0 | 0 | 0.2187 | 0.2036
0.074 | 0.01476 | 0.1599 | 0.04153 | 0 | 0 | 0.2162 | 0.2014
0.075 | 0.01441 | 0.1596 | 0.03981 | 0 | 0 | 0.2138 | 0.1994
0.076 | 0.01406 | 0.1592 | 0.03817 | 0 | 0 | 0.2115 | 0.1974
0.077 | 0.01373 | 0.1589 | 0.03663 | 0 | 0 | 0.2093 | 0.1955
0.078 | 0.01341 | 0.1586 | 0.03516 | 0 | 0 | 0.2072 | 0.1938
0.079 | 0.0131 | 0.1583 | 0.03377 | 0 | 0 | 0.2051 | 0.192
0.08 | 0.0128 | 0.158 | 0.03246 | 0 | 0 | 0.2032 | 0.1904
0.081 | 0.01251 | 0.1576 | 0.03121 | 0 | 0 | 0.2013 | 0.1888
0.082 | 0.01223 | 0.1573 | 0.03002 | 0 | 0 | 0.1996 | 0.1873
0.083 | 0.01196 | 0.157 | 0.02889 | 0 | 0 | 0.1978 | 0.1859
0.084 | 0.0117 | 0.1567 | 0.02782 | 0 | 0 | 0.1962 | 0.1845
0.085 | 0.01144 | 0.1563 | 0.02679 | 0 | 0 | 0.1946 | 0.1831
0.086 | 0.0112 | 0.156 | 0.02582 | 0 | 0 | 0.193 | 0.1818
0.087 | 0.01096 | 0.1557 | 0.02489 | 0 | 0 | 0.1916 | 0.1806
0.088 | 0.01073 | 0.1554 | 0.02401 | 0 | 0 | 0.1901 | 0.1794
0.089 | 0.01051 | 0.1551 | 0.02316 | 0 | 0 | 0.1887 | 0.1782
0.09 | 0.01029 | 0.1548 | 0.02236 | 0 | 0 | 0.1874 | 0.1771
0.091 | 0.01008 | 0.1544 | 0.02159 | 0 | 0 | 0.1861 | 0.176
0.092 | 0.009876 | 0.1541 | 0.02086 | 0 | 0 | 0.1848 | 0.175
0.093 | 0.009679 | 0.1538 | 0.02015 | 0 | 0 | 0.1836 | 0.174
0.094 | 0.009488 | 0.1535 | 0.01948 | 0 | 0 | 0.1825 | 0.173
0.095 | 0.009302 | 0.1532 | 0.01884 | 0 | 0 | 0.1813 | 0.172
0.096 | 0.009121 | 0.1529 | 0.01823 | 0 | 0 | 0.1802 | 0.1711
0.097 | 0.008946 | 0.1525 | 0.01764 | 0 | 0 | 0.1791 | 0.1702
0.098 | 0.008775 | 0.1522 | 0.01708 | 0 | 0 | 0.1781 | 0.1693
0.099 | 0.00861 | 0.1519 | 0.01654 | 0 | 0 | 0.1771 | 0.1685
0.1 | 0.008449 | 0.1516 | 0.01602 | 0 | 0 | 0.1761 | 0.1676
0.101 | 0.008292 | 0.1513 | 0.01552 | 0 | 0 | 0.1751 | 0.1668
0.102 | 0.00814 | 0.151 | 0.01504 | 0 | 0 | 0.1742 | 0.166
0.103 | 0.007992 | 0.1507 | 0.01459 | 0 | 0 | 0.1733 | 0.1653
0.104 | 0.007848 | 0.1504 | 0.01415 | 0 | 0 | 0.1724 | 0.1645
0.105 | 0.007708 | 0.1501 | 0.01372 | 0 | 0 | 0.1715 | 0.1638
0.106 | 0.007571 | 0.1498 | 0.01332 | 0 | 0 | 0.1707 | 0.1631
0.107 | 0.007438 | 0.1495 | 0.01293 | 0 | 0 | 0.1699 | 0.1624
0.108 | 0.007309 | 0.1492 | 0.01255 | 0 | 0 | 0.1691 | 0.1617
0.109 | 0.007183 | 0.1489 | 0.01219 | 0 | 0 | 0.1683 | 0.1611
0.11 | 0.00706 | 0.1486 | 0.01184 | 0 | 0 | 0.1675 | 0.1604
0.111 | 0.00694 | 0.1483 | 0.01151 | 0 | 0 | 0.1667 | 0.1598
0.112 | 0.006823 | 0.148 | 0.01119 | 0 | 0 | 0.166 | 0.1592
0.113 | 0.006709 | 0.1477 | 0.01088 | 0 | 0 | 0.1653 | 0.1586
0.114 | 0.006598 | 0.1474 | 0.01058 | 0 | 0 | 0.1646 | 0.158
0.115 | 0.00649 | 0.1471 | 0.01029 | 0 | 0 | 0.1639 | 0.1574
0.116 | 0.006384 | 0.1468 | 0.01001 | 0 | 0 | 0.1632 | 0.1568
0.117 | 0.006281 | 0.1465 | 0.009744 | 0 | 0 | 0.1626 | 0.1563
0.118 | 0.00618 | 0.1463 | 0.009485 | 0 | 0 | 0.1619 | 0.1557
0.119 | 0.006082 | 0.146 | 0.009235 | 0 | 0 | 0.1613 | 0.1552
0.12 | 0.005986 | 0.1457 | 0.008994 | 0 | 0 | 0.1607 | 0.1547
0.121 | 0.005892 | 0.1454 | 0.008761 | 0 | 0 | 0.1601 | 0.1542
0.122 | 0.005801 | 0.1451 | 0.008535 | 0 | 0 | 0.1595 | 0.1537
0.123 | 0.005711 | 0.1448 | 0.008318 | 0 | 0 | 0.1589 | 0.1532
0.124 | 0.005624 | 0.1446 | 0.008108 | 0 | 0 | 0.1583 | 0.1527
0.125 | 0.005538 | 0.1443 | 0.007904 | 0 | 0 | 0.1577 | 0.1522
0.126 | 0.005455 | 0.144 | 0.007708 | 0 | 0 | 0.1572 | 0.1517
0.127 | 0.005373 | 0.1437 | 0.007518 | 0 | 0 | 0.1566 | 0.1512
0.128 | 0.005293 | 0.1435 | 0.007334 | 0 | 0 | 0.1561 | 0.1508
0.129 | 0.005215 | 0.1432 | 0.007155 | 0 | 0 | 0.1555 | 0.1503
0.13 | 0.005139 | 0.1429 | 0.006983 | 0 | 0 | 0.155 | 0.1499
0.131 | 0.005064 | 0.1426 | 0.006816 | 0 | 0 | 0.1545 | 0.1494
0.132 | 0.004991 | 0.1424 | 0.006654 | 0 | 0 | 0.154 | 0.149
0.133 | 0.004919 | 0.1421 | 0.006498 | 0 | 0 | 0.1535 | 0.1486
0.134 | 0.004849 | 0.1418 | 0.006346 | 0 | 0 | 0.153 | 0.1482
0.135 | 0.004781 | 0.1416 | 0.006198 | 0 | 0 | 0.1525 | 0.1478
0.136 | 0.004714 | 0.1413 | 0.006056 | 0 | 0 | 0.1521 | 0.1474
0.137 | 0.004648 | 0.141 | 0.005917 | 0 | 0 | 0.1516 | 0.1469
0.138 | 0.004584 | 0.1408 | 0.005783 | 0 | 0 | 0.1511 | 0.1466
0.139 | 0.004521 | 0.1405 | 0.005653 | 0 | 0 | 0.1507 | 0.1462
0.14 | 0.004459 | 0.1402 | 0.005527 | 0 | 0 | 0.1502 | 0.1458
0.145 | 0.004168 | 0.139 | 0.004948 | 0 | 0 | 0.1481 | 0.1439
0.15 | 0.003905 | 0.1377 | 0.004447 | 0 | 0 | 0.1461 | 0.1422
0.155 | 0.003666 | 0.1365 | 0.004011 | 0 | 0 | 0.1442 | 0.1405
0.16 | 0.003448 | 0.1353 | 0.00363 | 0 | 0 | 0.1424 | 0.1389
0.165 | 0.003249 | 0.1341 | 0.003296 | 0 | 0 | 0.1407 | 0.1374
0.17 | 0.003067 | 0.133 | 0.003002 | 0 | 0 | 0.139 | 0.136
0.175 | 0.002899 | 0.1319 | 0.002742 | 0 | 0 | 0.1375 | 0.1346
0.18 | 0.002745 | 0.1308 | 0.002511 | 0 | 0 | 0.136 | 0.1333
0.185 | 0.002603 | 0.1297 | 0.002305 | 0 | 0 | 0.1346 | 0.132
0.19 | 0.002471 | 0.1287 | 0.002122 | 0 | 0 | 0.1332 | 0.1308
0.195 | 0.002349 | 0.1276 | 0.001957 | 0 | 0 | 0.1319 | 0.1296
0.2 | 0.002236 | 0.1266 | 0.001809 | 0 | 0 | 0.1307 | 0.1284
0.21 | 0.002033 | 0.1247 | 0.001556 | 0 | 0 | 0.1283 | 0.1263
0.22 | 0.001857 | 0.1229 | 0.001348 | 0 | 0 | 0.1261 | 0.1242
0.23 | 0.001702 | 0.1211 | 0.001176 | 0 | 0 | 0.124 | 0.1223
0.24 | 0.001566 | 0.1194 | 0.001033 | 0 | 0 | 0.122 | 0.1204
0.25 | 0.001445 | 0.1178 | 0.0009121 | 0 | 0 | 0.1201 | 0.1187
0.26 | 0.001338 | 0.1162 | 0.00081 | 0 | 0 | 0.1183 | 0.117
0.28 | 0.001157 | 0.1132 | 0.0006483 | 0 | 0 | 0.115 | 0.1139
0.3 | 0.00101 | 0.1105 | 0.000528 | 0 | 0 | 0.112 | 0.111
0.32 | 0.0008888 | 0.1079 | 0.0004366 | 0 | 0 | 0.1092 | 0.1083
0.34 | 0.0007884 | 0.1055 | 0.0003659 | 0 | 0 | 0.1066 | 0.1058
0.36 | 0.0007041 | 0.1032 | 0.0003104 | 0 | 0 | 0.1042 | 0.1035
0.38 | 0.0006326 | 0.1011 | 0.000266 | 0 | 0 | 0.102 | 0.1013
0.4 | 0.0005715 | 0.09908 | 0.0002302 | 0 | 0 | 0.09988 | 0.09931
0.42 | 0.0005187 | 0.09719 | 0.0002009 | 0 | 0 | 0.09791 | 0.09739
0.44 | 0.000473 | 0.0954 | 0.0001768 | 0 | 0 | 0.09605 | 0.09558
0.46 | 0.000433 | 0.09371 | 0.0001566 | 0 | 0 | 0.0943 | 0.09386
0.48 | 0.0003979 | 0.09209 | 0.0001397 | 0 | 0 | 0.09263 | 0.09223
0.5 | 0.0003668 | 0.09055 | 0.0001254 | 0 | 0 | 0.09104 | 0.09067
0.55 | 0.0003034 | 0.08698 | 9.798e-05 | 0 | 0 | 0.08738 | 0.08708
0.6 | 0.0002552 | 0.08378 | 7.858e-05 | 0 | 0 | 0.08412 | 0.08386
0.65 | 0.0002176 | 0.08089 | 6.432e-05 | 0 | 0 | 0.08118 | 0.08096
0.7 | 0.0001877 | 0.07827 | 5.374e-05 | 0 | 0 | 0.07851 | 0.07832
0.75 | 0.0001635 | 0.07585 | 4.583e-05 | 0 | 0 | 0.07606 | 0.0759
0.8 | 0.0001438 | 0.07363 | 3.987e-05 | 0 | 0 | 0.07381 | 0.07367
0.85 | 0.0001274 | 0.07156 | 3.53e-05 | 0 | 0 | 0.07173 | 0.0716
0.9 | 0.0001136 | 0.06965 | 3.148e-05 | 0 | 0 | 0.06979 | 0.06968
0.95 | 0.000102 | 0.06787 | 2.804e-05 | 0 | 0 | 0.068 | 0.0679
1 | 9.209e-05 | 0.06622 | 2.479e-05 | 0 | 0 | 0.06634 | 0.06625
2 | 2.305e-05 | 0.04594 | 7.124e-06 | 0.0004886 | 0 | 0.04646 | 0.04643
3 | 1.025e-05 | 0.03614 | 3.909e-06 | 0.001388 | 1.264e-05 | 0.03755 | 0.03754
4 | 5.765e-06 | 0.03015 | 2.651e-06 | 0.002244 | 5.163e-05 | 0.03245 | 0.03245
5 | 3.689e-06 | 0.02603 | 1.995e-06 | 0.003005 | 0.0001029 | 0.02915 | 0.02914
6 | 2.562e-06 | 0.02301 | 1.595e-06 | 0.003688 | 0.000158 | 0.02686 | 0.02686
7 | 1.882e-06 | 0.02068 | 1.326e-06 | 0.004296 | 0.000213 | 0.0252 | 0.02519
8 | 1.441e-06 | 0.01882 | 1.134e-06 | 0.004844 | 0.0002664 | 0.02394 | 0.02393
9 | 1.138e-06 | 0.0173 | 9.904e-07 | 0.005341 | 0.0003177 | 0.02296 | 0.02296
10 | 9.223e-07 | 0.01603 | 8.785e-07 | 0.005791 | 0.0003664 | 0.02219 | 0.02219
11 | 7.622e-07 | 0.01495 | 7.893e-07 | 0.006201 | 0.0004128 | 0.02157 | 0.02157
12 | 6.405e-07 | 0.01402 | 7.164e-07 | 0.006577 | 0.0004566 | 0.02106 | 0.02106
13 | 5.458e-07 | 0.01321 | 6.557e-07 | 0.006926 | 0.000498 | 0.02063 | 0.02063
14 | 4.704e-07 | 0.0125 | 6.043e-07 | 0.007252 | 0.0005375 | 0.02029 | 0.02029
15 | 4.099e-07 | 0.01187 | 5.606e-07 | 0.007554 | 0.0005749 | 0.02 | 0.01999
16 | 3.602e-07 | 0.0113 | 5.227e-07 | 0.007841 | 0.0006107 | 0.01975 | 0.01975
17 | 3.191e-07 | 0.01079 | 4.894e-07 | 0.008111 | 0.0006447 | 0.01955 | 0.01954
18 | 2.846e-07 | 0.01033 | 4.601e-07 | 0.008364 | 0.0006772 | 0.01937 | 0.01937
19 | 2.554e-07 | 0.009906 | 4.342e-07 | 0.008603 | 0.0007082 | 0.01922 | 0.01922
20 | 2.305e-07 | 0.009523 | 4.11e-07 | 0.00883 | 0.000738 | 0.01909 | 0.01909

## Teflon
Energy (MeV) | Coherent (cm2/g) | Incoherent (cm2/g) | Photoelectric (cm2/g) | Nuclear Pair (cm2/g) | Electron Pair (cm2/g) | Total With Coherent (cm2/g) | Total Without Coherent (cm2/g)
0.02 | 0.1025 | 0.1473 | 0.717 | 0 | 0 | 0.9668 | 0.8643
0.021 | 0.09531 | 0.1485 | 0.6131 | 0 | 0 | 0.8569 | 0.7616
0.022 | 0.08883 | 0.1495 | 0.528 | 0 | 0 | 0.7663 | 0.6775
0.023 | 0.08298 | 0.1504 | 0.4576 | 0 | 0 | 0.691 | 0.608
0.024 | 0.07769 | 0.1512 | 0.399 | 0 | 0 | 0.6279 | 0.5502
0.025 | 0.07288 | 0.152 | 0.3497 | 0 | 0 | 0.5745 | 0.5017
0.026 | 0.0685 | 0.1526 | 0.3081 | 0 | 0 | 0.5292 | 0.4607
0.027 | 0.06449 | 0.1531 | 0.2727 | 0 | 0 | 0.4903 | 0.4258
0.028 | 0.06082 | 0.1536 | 0.2424 | 0 | 0 | 0.4568 | 0.396
0.029 | 0.05745 | 0.154 | 0.2163 | 0 | 0 | 0.4278 | 0.3703
0.03 | 0.05435 | 0.1544 | 0.1938 | 0 | 0 | 0.4025 | 0.3481
0.031 | 0.05148 | 0.1547 | 0.1742 | 0 | 0 | 0.3803 | 0.3289
0.032 | 0.04883 | 0.1549 | 0.1571 | 0 | 0 | 0.3609 | 0.312
0.033 | 0.04638 | 0.1551 | 0.1421 | 0 | 0 | 0.3436 | 0.2973
0.034 | 0.0441 | 0.1553 | 0.129 | 0 | 0 | 0.3283 | 0.2842
0.035 | 0.04198 | 0.1554 | 0.1174 | 0 | 0 | 0.3147 | 0.2727
0.036 | 0.04001 | 0.1555 | 0.1071 | 0 | 0 | 0.3025 | 0.2625
0.037 | 0.03818 | 0.1555 | 0.09791 | 0 | 0 | 0.2916 | 0.2534
0.038 | 0.03646 | 0.1555 | 0.08975 | 0 | 0 | 0.2817 | 0.2453
0.039 | 0.03486 | 0.1555 | 0.08245 | 0 | 0 | 0.2728 | 0.238
0.04 | 0.03336 | 0.1555 | 0.07591 | 0 | 0 | 0.2648 | 0.2314
0.041 | 0.03195 | 0.1554 | 0.07002 | 0 | 0 | 0.2574 | 0.2255
0.042 | 0.03063 | 0.1554 | 0.06472 | 0 | 0 | 0.2507 | 0.2201
0.043 | 0.02939 | 0.1553 | 0.05993 | 0 | 0 | 0.2446 | 0.2152
0.044 | 0.02822 | 0.1551 | 0.05558 | 0 | 0 | 0.2389 | 0.2107
0.045 | 0.02712 | 0.155 | 0.05164 | 0 | 0 | 0.2338 | 0.2067
0.046 | 0.02609 | 0.1549 | 0.04806 | 0 | 0 | 0.229 | 0.2029
0.047 | 0.02511 | 0.1547 | 0.04479 | 0 | 0 | 0.2246 | 0.1995
0.048 | 0.02418 | 0.1545 | 0.0418 | 0 | 0 | 0.2205 | 0.1963
0.049 | 0.02331 | 0.1543 | 0.03907 | 0 | 0 | 0.2167 | 0.1934
0.05 | 0.02248 | 0.1541 | 0.03656 | 0 | 0 | 0.2132 | 0.1907
0.051 | 0.02169 | 0.1539 | 0.03426 | 0 | 0 | 0.2099 | 0.1882
0.052 | 0.02095 | 0.1537 | 0.03214 | 0 | 0 | 0.2068 | 0.1859
0.053 | 0.02024 | 0.1535 | 0.03019 | 0 | 0 | 0.2039 | 0.1837
0.054 | 0.01957 | 0.1533 | 0.02839 | 0 | 0 | 0.2012 | 0.1816
0.055 | 0.01893 | 0.153 | 0.02673 | 0 | 0 | 0.1987 | 0.1797
0.056 | 0.01832 | 0.1528 | 0.02519 | 0 | 0 | 0.1963 | 0.178
0.057 | 0.01774 | 0.1525 | 0.02376 | 0 | 0 | 0.194 | 0.1763
0.058 | 0.01719 | 0.1523 | 0.02244 | 0 | 0 | 0.1919 | 0.1747
0.059 | 0.01666 | 0.152 | 0.02122 | 0 | 0 | 0.1899 | 0.1732
0.06 | 0.01616 | 0.1517 | 0.02007 | 0 | 0 | 0.188 | 0.1718
0.061 | 0.01568 | 0.1515 | 0.01901 | 0 | 0 | 0.1862 | 0.1705
0.062 | 0.01522 | 0.1512 | 0.01802 | 0 | 0 | 0.1844 | 0.1692
0.063 | 0.01478 | 0.1509 | 0.0171 | 0 | 0 | 0.1828 | 0.168
0.064 | 0.01436 | 0.1506 | 0.01623 | 0 | 0 | 0.1812 | 0.1669
0.065 | 0.01395 | 0.1504 | 0.01542 | 0 | 0 | 0.1797 | 0.1658
0.066 | 0.01357 | 0.1501 | 0.01467 | 0 | 0 | 0.1783 | 0.1647
0.067 | 0.0132 | 0.1498 | 0.01396 | 0 | 0 | 0.1769 | 0.1637
0.068 | 0.01284 | 0.1495 | 0.0133 | 0 | 0 | 0.1756 | 0.1628
0.069 | 0.0125 | 0.1492 | 0.01267 | 0 | 0 | 0.1744 | 0.1619
0.07 | 0.01217 | 0.1489 | 0.01209 | 0 | 0 | 0.1732 | 0.161
0.071 | 0.01185 | 0.1486 | 0.01153 | 0 | 0 | 0.172 | 0.1602
0.072 | 0.01155 | 0.1483 | 0.01102 | 0 | 0 | 0.1709 | 0.1593
0.073 | 0.01126 | 0.148 | 0.01053 | 0 | 0 | 0.1698 | 0.1586
0.074 | 0.01098 | 0.1477 | 0.01007 | 0 | 0 | 0.1688 | 0.1578
0.075 | 0.0107 | 0.1474 | 0.009631 | 0 | 0 | 0.1678 | 0.1571
0.076 | 0.01044 | 0.1471 | 0.00922 | 0 | 0 | 0.1668 | 0.1564
0.077 | 0.01019 | 0.1468 | 0.008832 | 0 | 0 | 0.1659 | 0.1557
0.078 | 0.009949 | 0.1465 | 0.008464 | 0 | 0 | 0.165 | 0.155
0.079 | 0.009714 | 0.1463 | 0.008117 | 0 | 0 | 0.1641 | 0.1544
0.08 | 0.009488 | 0.146 | 0.007787 | 0 | 0 | 0.1632 | 0.1537
0.081 | 0.00927 | 0.1457 | 0.007475 | 0 | 0 | 0.1624 | 0.1531
0.082 | 0.009059 | 0.1454 | 0.00718 | 0 | 0 | 0.1616 | 0.1525
0.083 | 0.008854 | 0.1451 | 0.006899 | 0 | 0 | 0.1608 | 0.152
0.084 | 0.008657 | 0.1448 | 0.006632 | 0 | 0 | 0.1601 | 0.1514
0.085 | 0.008466 | 0.1445 | 0.006379 | 0 | 0 | 0.1593 | 0.1508
0.086 | 0.008282 | 0.1442 | 0.006138 | 0 | 0 | 0.1586 | 0.1503
0.087 | 0.008103 | 0.1439 | 0.005909 | 0 | 0 | 0.1579 | 0.1498
0.088 | 0.00793 | 0.1436 | 0.005691 | 0 | 0 | 0.1572 | 0.1493
0.089 | 0.007762 | 0.1433 | 0.005483 | 0 | 0 | 0.1565 | 0.1488
0.09 | 0.0076 | 0.143 | 0.005285 | 0 | 0 | 0.1559 | 0.1483
0.091 | 0.007442 | 0.1427 | 0.005096 | 0 | 0 | 0.1552 | 0.1478
0.092 | 0.00729 | 0.1424 | 0.004916 | 0 | 0 | 0.1546 | 0.1473
0.093 | 0.007142 | 0.1421 | 0.004745 | 0 | 0 | 0.154 | 0.1469
0.094 | 0.006998 | 0.1418 | 0.004581 | 0 | 0 | 0.1534 | 0.1464
0.095 | 0.006859 | 0.1415 | 0.004424 | 0 | 0 | 0.1528 | 0.1459
0.096 | 0.006724 | 0.1412 | 0.004274 | 0 | 0 | 0.1522 | 0.1455
0.097 | 0.006593 | 0.1409 | 0.004131 | 0 | 0 | 0.1517 | 0.1451
0.098 | 0.006465 | 0.1407 | 0.003994 | 0 | 0 | 0.1511 | 0.1446
0.099 | 0.006341 | 0.1404 | 0.003863 | 0 | 0 | 0.1506 | 0.1442
0.1 | 0.006221 | 0.1401 | 0.003738 | 0 | 0 | 0.15 | 0.1438
0.101 | 0.006104 | 0.1398 | 0.003617 | 0 | 0 | 0.1495 | 0.1434
0.102 | 0.00599 | 0.1395 | 0.003502 | 0 | 0 | 0.149 | 0.143
0.103 | 0.00588 | 0.1392 | 0.003392 | 0 | 0 | 0.1485 | 0.1426
0.104 | 0.005772 | 0.1389 | 0.003286 | 0 | 0 | 0.148 | 0.1422
0.105 | 0.005667 | 0.1387 | 0.003184 | 0 | 0 | 0.1475 | 0.1418
0.106 | 0.005565 | 0.1384 | 0.003087 | 0 | 0 | 0.147 | 0.1415
0.107 | 0.005466 | 0.1381 | 0.002993 | 0 | 0 | 0.1466 | 0.1411
0.108 | 0.00537 | 0.1378 | 0.002903 | 0 | 0 | 0.1461 | 0.1407
0.109 | 0.005276 | 0.1375 | 0.002817 | 0 | 0 | 0.1456 | 0.1404
0.11 | 0.005184 | 0.1373 | 0.002734 | 0 | 0 | 0.1452 | 0.14
0.111 | 0.005095 | 0.137 | 0.002654 | 0 | 0 | 0.1447 | 0.1396
0.112 | 0.005008 | 0.1367 | 0.002577 | 0 | 0 | 0.1443 | 0.1393
0.113 | 0.004923 | 0.1364 | 0.002503 | 0 | 0 | 0.1439 | 0.1389
0.114 | 0.00484 | 0.1362 | 0.002432 | 0 | 0 | 0.1434 | 0.1386
0.115 | 0.00476 | 0.1359 | 0.002363 | 0 | 0 | 0.143 | 0.1383
0.116 | 0.004681 | 0.1356 | 0.002297 | 0 | 0 | 0.1426 | 0.1379
0.117 | 0.004604 | 0.1354 | 0.002233 | 0 | 0 | 0.1422 | 0.1376
0.118 | 0.00453 | 0.1351 | 0.002172 | 0 | 0 | 0.1418 | 0.1373
0.119 | 0.004457 | 0.1348 | 0.002113 | 0 | 0 | 0.1414 | 0.1369
0.12 | 0.004385 | 0.1346 | 0.002056 | 0 | 0 | 0.141 | 0.1366
0.121 | 0.004316 | 0.1343 | 0.002001 | 0 | 0 | 0.1406 | 0.1363
0.122 | 0.004248 | 0.134 | 0.001948 | 0 | 0 | 0.1402 | 0.136
0.123 | 0.004181 | 0.1338 | 0.001896 | 0 | 0 | 0.1398 | 0.1357
0.124 | 0.004116 | 0.1335 | 0.001847 | 0 | 0 | 0.1395 | 0.1353
0.125 | 0.004053 | 0.1332 | 0.001799 | 0 | 0 | 0.1391 | 0.135
0.126 | 0.003991 | 0.133 | 0.001753 | 0 | 0 | 0.1387 | 0.1347
0.127 | 0.00393 | 0.1327 | 0.001708 | 0 | 0 | 0.1384 | 0.1344
0.128 | 0.003871 | 0.1325 | 0.001665 | 0 | 0 | 0.138 | 0.1341
0.129 | 0.003813 | 0.1322 | 0.001623 | 0 | 0 | 0.1377 | 0.1338
0.13 | 0.003757 | 0.132 | 0.001583 | 0 | 0 | 0.1373 | 0.1335
0.131 | 0.003701 | 0.1317 | 0.001544 | 0 | 0 | 0.137 | 0.1333
0.132 | 0.003647 | 0.1315 | 0.001506 | 0 | 0 | 0.1366 | 0.133
0.133 | 0.003594 | 0.1312 | 0.001469 | 0 | 0 | 0.1363 | 0.1327
0.134 | 0.003542 | 0.131 | 0.001434 | 0 | 0 | 0.1359 | 0.1324
0.135 | 0.003492 | 0.1307 | 0.001399 | 0 | 0 | 0.1356 | 0.1321
0.136 | 0.003442 | 0.1305 | 0.001366 | 0 | 0 | 0.1353 | 0.1318
0.137 | 0.003393 | 0.1302 | 0.001334 | 0 | 0 | 0.1349 | 0.1316
0.138 | 0.003346 | 0.13 | 0.001303 | 0 | 0 | 0.1346 | 0.1313
0.139 | 0.003299 | 0.1297 | 0.001272 | 0 | 0 | 0.1343 | 0.131
0.14 | 0.003254 | 0.1295 | 0.001243 | 0 | 0 | 0.134 | 0.1307
0.145 | 0.003039 | 0.1283 | 0.001109 | 0 | 0 | 0.1324 | 0.1294
0.15 | 0.002844 | 0.1271 | 0.0009932 | 0 | 0 | 0.131 | 0.1281
0.155 | 0.002668 | 0.126 | 0.000893 | 0 | 0 | 0.1296 | 0.1269
0.16 | 0.002507 | 0.1249 | 0.0008057 | 0 | 0 | 0.1282 | 0.1257
0.165 | 0.002361 | 0.1238 | 0.0007294 | 0 | 0 | 0.1269 | 0.1245
0.17 | 0.002227 | 0.1227 | 0.0006623 | 0 | 0 | 0.1256 | 0.1234
0.175 | 0.002103 | 0.1217 | 0.0006033 | 0 | 0 | 0.1244 | 0.1223
0.18 | 0.00199 | 0.1207 | 0.000551 | 0 | 0 | 0.1232 | 0.1212
0.185 | 0.001886 | 0.1197 | 0.0005046 | 0 | 0 | 0.1221 | 0.1202
0.19 | 0.001789 | 0.1187 | 0.0004633 | 0 | 0 | 0.121 | 0.1192
0.195 | 0.0017 | 0.1178 | 0.0004264 | 0 | 0 | 0.1199 | 0.1182
0.2 | 0.001617 | 0.1169 | 0.0003933 | 0 | 0 | 0.1189 | 0.1173
0.21 | 0.001469 | 0.1151 | 0.0003368 | 0 | 0 | 0.1169 | 0.1154
0.22 | 0.00134 | 0.1134 | 0.0002908 | 0 | 0 | 0.115 | 0.1137
0.23 | 0.001227 | 0.1117 | 0.0002528 | 0 | 0 | 0.1132 | 0.112
0.24 | 0.001128 | 0.1102 | 0.0002213 | 0 | 0 | 0.1115 | 0.1104
0.25 | 0.001041 | 0.1086 | 0.0001949 | 0 | 0 | 0.1099 | 0.1088
0.26 | 0.000963 | 0.1072 | 0.0001726 | 0 | 0 | 0.1083 | 0.1074
0.28 | 0.0008314 | 0.1045 | 0.0001374 | 0 | 0 | 0.1054 | 0.1046
0.3 | 0.000725 | 0.1019 | 0.0001114 | 0 | 0 | 0.1027 | 0.102
0.32 | 0.0006378 | 0.09953 | 9.18e-05 | 0 | 0 | 0.1003 | 0.09962
0.34 | 0.0005654 | 0.09731 | 7.667e-05 | 0 | 0 | 0.09795 | 0.09738
0.36 | 0.0005046 | 0.09522 | 6.482e-05 | 0 | 0 | 0.09579 | 0.09529
0.38 | 0.0004531 | 0.09326 | 5.54e-05 | 0 | 0 | 0.09377 | 0.09332
0.4 | 0.0004091 | 0.09141 | 4.78e-05 | 0 | 0 | 0.09187 | 0.09146
0.42 | 0.0003712 | 0.08967 | 4.161e-05 | 0 | 0 | 0.09008 | 0.08971
0.44 | 0.0003383 | 0.08801 | 3.652e-05 | 0 | 0 | 0.08838 | 0.08805
0.46 | 0.0003096 | 0.08644 | 3.229e-05 | 0 | 0 | 0.08678 | 0.08647
0.48 | 0.0002844 | 0.08494 | 2.876e-05 | 0 | 0 | 0.08525 | 0.08497
0.5 | 0.0002621 | 0.08351 | 2.578e-05 | 0 | 0 | 0.0838 | 0.08354
0.55 | 0.0002167 | 0.08022 | 2.009e-05 | 0 | 0 | 0.08046 | 0.08024
0.6 | 0.0001822 | 0.07728 | 1.606e-05 | 0 | 0 | 0.07747 | 0.07729
0.65 | 0.0001553 | 0.07462 | 1.308e-05 | 0 | 0 | 0.07478 | 0.07463
0.7 | 0.0001339 | 0.07219 | 1.088e-05 | 0 | 0 | 0.07234 | 0.0722
0.75 | 0.0001167 | 0.06997 | 9.271e-06 | 0 | 0 | 0.0701 | 0.06998
0.8 | 0.0001026 | 0.06792 | 8.109e-06 | 0 | 0 | 0.06803 | 0.06793
0.85 | 9.086e-05 | 0.06601 | 7.251e-06 | 0 | 0 | 0.06611 | 0.06602
0.9 | 8.106e-05 | 0.06425 | 6.516e-06 | 0 | 0 | 0.06433 | 0.06425
0.95 | 7.276e-05 | 0.0626 | 5.793e-06 | 0 | 0 | 0.06268 | 0.06261
1 | 6.566e-05 | 0.06108 | 5.041e-06 | 0 | 0 | 0.06115 | 0.06108
2 | 1.642e-05 | 0.04236 | 1.457e-06 | 0.0004219 | 0 | 0.0428 | 0.04278
3 | 7.3e-06 | 0.03333 | 8.126e-07 | 0.001207 | 1.166e-05 | 0.03456 | 0.03455
4 | 4.107e-06 | 0.0278 | 5.572e-07 | 0.001958 | 4.762e-05 | 0.02981 | 0.02981
5 | 2.628e-06 | 0.02401 | 4.222e-07 | 0.002627 | 9.487e-05 | 0.02674 | 0.02673
6 | 1.825e-06 | 0.02122 | 3.394e-07 | 0.003227 | 0.0001457 | 0.0246 | 0.0246
7 | 1.341e-06 | 0.01908 | 2.833e-07 | 0.003763 | 0.0001965 | 0.02304 | 0.02303
8 | 1.027e-06 | 0.01736 | 2.43e-07 | 0.004243 | 0.0002458 | 0.02185 | 0.02185
9 | 8.11e-07 | 0.01596 | 2.127e-07 | 0.00468 | 0.000293 | 0.02093 | 0.02093
10 | 6.571e-07 | 0.01478 | 1.891e-07 | 0.005078 | 0.0003381 | 0.0202 | 0.0202
11 | 5.43e-07 | 0.01379 | 1.702e-07 | 0.00544 | 0.0003809 | 0.01961 | 0.01961
12 | 4.561e-07 | 0.01293 | 1.547e-07 | 0.005771 | 0.0004211 | 0.01912 | 0.01912
13 | 3.888e-07 | 0.01218 | 1.418e-07 | 0.006078 | 0.0004593 | 0.01872 | 0.01872
14 | 3.353e-07 | 0.01153 | 1.308e-07 | 0.006364 | 0.0004959 | 0.01839 | 0.01839
15 | 2.921e-07 | 0.01094 | 1.214e-07 | 0.006632 | 0.0005306 | 0.01811 | 0.01811
16 | 2.566e-07 | 0.01042 | 1.133e-07 | 0.006883 | 0.0005633 | 0.01787 | 0.01787
17 | 2.273e-07 | 0.00995 | 1.062e-07 | 0.007121 | 0.0005947 | 0.01767 | 0.01767
18 | 2.028e-07 | 0.009524 | 9.993e-08 | 0.007346 | 0.0006248 | 0.0175 | 0.0175
19 | 1.82e-07 | 0.009137 | 9.435e-08 | 0.007558 | 0.0006535 | 0.01735 | 0.01735
20 | 1.642e-07 | 0.008783 | 8.936e-08 | 0.007759 | 0.000681 | 0.01722 | 0.01722

## Iodine
Energy (MeV) | Total With Coherent (cm2/g)
0.02 | 25.4
0.021 | 22.284
0.022 | 19.6698
0.023 | 17.4588
0.024 | 15.5751
0.025 | 13.9597
0.026 | 12.5656
0.027 | 11.3558
0.028 | 10.3003
0.029 | 9.3749
0.03 | 8.56
0.031 | 7.8438
0.032 | 7.2074
0.033 | 6.64
0.034 | 33.5904
0.035 | 31.1735
0.036 | 28.9913
0.037 | 27.0156
0.038 | 25.2219
0.039 | 23.5894
0.04 | 22.1
0.041 | 20.7124
0.042 | 19.4424
0.043 | 18.2773
0.044 | 17.2066
0.045 | 16.2205
0.046 | 15.3108
0.047 | 14.4701
0.048 | 13.6918
0.049 | 12.9702
0.05 | 12.3
0.051 | 11.67
0.052 | 11.0836
0.053 | 10.537
0.054 | 10.0268
0.055 | 9.55
0.056 | 9.1039
0.057 | 8.6859
0.058 | 8.2939
0.059 | 7.9259
0.06 | 7.58
0.061 | 7.252
0.062 | 6.9432
0.063 | 6.6522
0.064 | 6.3776
0.065 | 6.1184
0.066 | 5.8735
0.067 | 5.6418
0.068 | 5.4225
0.069 | 5.2147
0.07 | 5.0177
0.071 | 4.8308
0.072 | 4.6533
0.073 | 4.4847
0.074 | 4.3243
0.075 | 4.1717
0.076 | 4.0265
0.077 | 3.888
0.078 | 3.7561
0.079 | 3.6302
0.08 | 3.51
0.081 | 3.396
0.082 | 3.2871
0.083 | 3.1829
0.084 | 3.0832
0.085 | 2.9878
0.086 | 2.8963
0.087 | 2.8087
0.088 | 2.7247
0.089 | 2.6441
0.09 | 2.5668
0.091 | 2.4925
0.092 | 2.4212
0.093 | 2.3526
0.094 | 2.2867
0.095 | 2.2233
0.096 | 2.1623
0.097 | 2.1035
0.098 | 2.047
0.099 | 1.9925
0.1 | 1.94
0.101 | 1.8919
0.102 | 1.8455
0.103 | 1.8007
0.104 | 1.7574
0.105 | 1.7155
0.106 | 1.675
0.107 | 1.6358
0.108 | 1.5979
0.109 | 1.5612
0.11 | 1.5256
0.111 | 1.4912
0.112 | 1.4579
0.113 | 1.4256
0.114 | 1.3942
0.115 | 1.3639
0.116 | 1.3344
0.117 | 1.3059
0.118 | 1.2781
0.119 | 1.2512
0.12 | 1.2251
0.121 | 1.1997
0.122 | 1.1751
0.123 | 1.1512
0.124 | 1.1279
0.125 | 1.1053
0.126
0.127
0.128
0.129
0.13
0.131
0.132
0.133
0.134
0.135
0.136
0.137
0.138
0.139
0.14
0.145
0.15
0.155
0.16
0.165
0.17
0.175
0.18
0.185
0.19
0.195
0.2
0.21
0.22
0.23
0.24
0.25
0.26
0.28
0.3
0.32
0.34
0.36
0.38
0.4
0.42
0.44
0.46
0.48
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20

## CorticalBone
Energy (MeV) | Total With Coherent (cm2/g)
0.02 | 4.001
0.021 | 3.5047
0.022 | 3.0889
0.023 | 2.7378
0.024 | 2.4391
0.025 | 2.1833
0.026 | 1.9628
0.027 | 1.7717
0.028 | 1.6051
0.029 | 1.4593
0.03 | 1.331
0.031 | 1.2299
0.032 | 1.1393
0.033 | 1.0579
0.034 | 0.98448
0.035 | 0.91807
0.036 | 0.85782
0.037 | 0.80302
0.038 | 0.75305
0.039 | 0.70736
0.04 | 0.6655
0.041 | 0.63315
0.042 | 0.60309
0.043 | 0.57512
0.044 | 0.54905
0.045 | 0.52471
0.046 | 0.50194
0.047 | 0.48062
0.048 | 0.46063
0.049 | 0.44185
0.05 | 0.4242
0.051 | 0.41068
0.052 | 0.39784
0.053 | 0.38563
0.054 | 0.37402
0.055 | 0.36296
0.056 | 0.35241
0.057 | 0.34236
0.058 | 0.33275
0.059 | 0.32358
0.06 | 0.3148
0.061 | 0.30862
0.062 | 0.30265
0.063 | 0.2969
0.064 | 0.29134
0.065 | 0.28597
0.066 | 0.28078
0.067 | 0.27576
0.068 | 0.2709
0.069 | 0.26619
0.07 | 0.26164
0.071 | 0.25722
0.072 | 0.25294
0.073 | 0.24879
0.074 | 0.24476
0.075 | 0.24085
0.076 | 0.23705
0.077 | 0.23336
0.078 | 0.22978
0.079 | 0.22629
0.08 | 0.2229
0.081 | 0.22063
0.082 | 0.21842
0.083 | 0.21625
0.084 | 0.21413
0.085 | 0.21205
0.086 | 0.21002
0.087 | 0.20803
0.088 | 0.20608
0.089 | 0.20417
0.09 | 0.20231
0.091 | 0.20047
0.092 | 0.19868
0.093 | 0.19692
0.094 | 0.19519
0.095 | 0.1935
0.096 | 0.19184
0.097 | 0.19021
0.098 | 0.18861
0.099 | 0.18704
0.1 | 0.1855
0.101 | 0.18447
0.102 | 0.18347
0.103 | 0.18247
0.104 | 0.18149
0.105 | 0.18053
0.106 | 0.17958
0.107 | 0.17864
0.108 | 0.17772
0.109 | 0.17681
0.11 | 0.17591
0.111 | 0.17502
0.112 | 0.17415
0.113 | 0.17329
0.114 | 0.17244
0.115 | 0.17161
0.116 | 0.17078
0.117 | 0.16997
0.118 | 0.16916
0.119 | 0.16837
0.12 | 0.16759
0.121 | 0.16681
0.122 | 0.16605
0.123 | 0.1653
0.124 | 0.16455
0.125 | 0.16382
0.126 | 0.16309
0.127 | 0.16238
0.128 | 0.16167
0.129 | 0.16097
0.13 | 0.16028
0.131 | 0.1596
0.132 | 0.15892
0.133 | 0.15826
0.134 | 0.1576
0.135 | 0.15695
0.136 | 0.1563
0.137 | 0.15567
0.138 | 0.15504
0.139 | 0.15441
0.14 | 0.1538
0.145
0.15
0.155
0.16
0.165
0.17
0.175
0.18
0.185
0.19
0.195
0.2
0.21
0.22
0.23
0.24
0.25
0.26
0.28
0.3
0.32
0.34
0.36
0.38
0.4
0.42
0.44
0.46
0.48
0.5
0.55
0.6
0.65
0.7
0.75
0.8
0.85
0.9
0.95
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20

## CaCl2
Energy (MeV) | Total With Coherent (cm2/g)
0.02 | 9.6565
0.021 | 8.3965
0.022 | 7.3486
0.023 | 6.4697
0.024 | 5.7268
0.025 | 5.0946
0.026 | 4.553
0.027 | 4.0863
0.028 | 3.6819
0.029 | 3.3297
0.03 | 3.0214
0.031 | 2.7618
0.032 | 2.5318
0.033 | 2.3271
0.034 | 2.1444
0.035 | 1.9807
0.036 | 1.8336
0.037 | 1.701
0.038 | 1.5812
0.039 | 1.4726
0.04 | 1.3739
0.041 | 1.2908
0.042 | 1.2146
0.043 | 1.1445
0.044 | 1.0799
0.045 | 1.0203
0.046 | 0.96517
0.047 | 0.91414
0.048 | 0.8668
0.049 | 0.82281
0.05 | 0.78189
0.051 | 0.74769
0.052 | 0.71562
0.053 | 0.68551
0.054 | 0.65719
0.055 | 0.63053
0.056 | 0.60541
0.057 | 0.58172
0.058 | 0.55934
0.059 | 0.53818
0.06 | 0.51817
0.061 | 0.50252
0.062 | 0.48759
0.063 | 0.47333
0.064 | 0.45971
0.065 | 0.44669
0.066 | 0.43423
0.067 | 0.4223
0.068 | 0.41087
0.069 | 0.39991
0.07 | 0.3894
0.071 | 0.3793
0.072 | 0.36961
0.073 | 0.3603
0.074 | 0.35135
0.075 | 0.34273
0.076 | 0.33444
0.077 | 0.32646
0.078 | 0.31876
0.079 | 0.31135
0.08 | 0.30419
0.081 | 0.29903
0.082 | 0.29401
0.083 | 0.28914
0.084 | 0.28441
0.085 | 0.27981
0.086 | 0.27534
0.087 | 0.27099
0.088 | 0.26676
0.089 | 0.26264
0.09 | 0.25864
0.091 | 0.25473
0.092 | 0.25093
0.093 | 0.24723
0.094 | 0.24362
0.095 | 0.24011
0.096 | 0.23668
0.097 | 0.23333
0.098 | 0.23007
0.099 | 0.22688
0.1 | 0.22377
0.101 | 0.22176
0.102 | 0.21978
0.103 | 0.21784
0.104 | 0.21594
0.105 | 0.21407
0.106 | 0.21224
0.107 | 0.21043
0.108 | 0.20867
0.109 | 0.20693
0.11 | 0.20522
0.111 | 0.20354
0.112 | 0.20189
0.113 | 0.20027
0.114 | 0.19868
0.115 | 0.19711
0.116 | 0.19557
0.117 | 0.19405
0.118 | 0.19256
0.119 | 0.19109
0.12 | 0.18965
0.121 | 0.18823
0.122 | 0.18683
0.123 | 0.18545
0.124 | 0.1841
0.125 | 0.18276
0.126 | 0.18145
0.127 | 0.18016
0.128 | 0.17888
0.129 | 0.17763
0.13 | 0.17639
0.131 | 0.17517
0.132 | 0.17397
0.133 | 0.17279
0.134 | 0.17162
0.135 | 0.17047
0.136 | 0.16934
0.137 | 0.16822
0.138 | 0.16712
0.139 | 0.16603
0.14 | 0.16496
0.141 | 0.1639
0.142 | 0.16286
0.143 | 0.16183
0.144 | 0.16081
0.145 | 0.15981
0.146 | 0.15882
0.147 | 0.15784
0.148 | 0.15688
0.149 | 0.15593
0.15 | 0.15499